# Blank Template

Generated from: DHF..pdf

============================================================

## BLANK TEMPLATE

This template was automatically generated from a filled document.

Please fill in the blank fields marked with underscores (_____).

PASSIM LIFESCIENCES LIMITED

Haryana- ______

DESIGN HISTORY FILE

Name of Product

Piston Nebulizer

Effective Date

27-09-2024

Document Number

PLL/DHF/001

Next Review Date

26-09-2029

Revision Number

R-01

Page Number

Page 1 of 9

Design History File

(Piston Nebulizer)

Manufacturing Site:

Passim Lifesciences Limited

Plot No. 45, Industrial Area Phase 2,

Haryana- ______

This design history file is prepared as per requirement of ISO 13485:2016, clause 7.3 design and

development requirement. It is maintained at above said manufacturing site.

Prepared By:

Fira

27-09-2024

MASTER COPY

MAST

Approved by:

27/09/2024

PASSIM LIFESCIENCES LIMITED

Haryana- ______

DESIGN HISTORY FILE

Name of Product

Piston Nebulizer

Effective Date

27-09-2024

Document Number

PLL/DHF/001

Next Review Date

26-09-2029

Revision Number

R-01

Page Number

Page 2 of 9

1. Design and Development Planning

Description of Project: ______________________________

______________________________

planned due to emerging need of active medical device and its accessories in market to fulfil

the need of user at affordable prices and for better business growth.

Design planning of Piston Nebulizer was initiated to add new product range also and

upcoming need to active medical device accessories.

Requirement was raised by the marketing team to see the product market growth.

Design and Development Planning Approach:

Following design and development stages applied to the product Piston Nebulizer :

Design & Development Planning +

Control of Design & Development Changes

Design & Development Inputs

Design & Development Output

Design & Development Review

Design & Development

Verification

Design & Development

Validation

•

•

The review(s) needed at each design and development stage

The verification, validation, and design transfer activities that are appropriate at each

design and development stage.

Define the responsibilities and authorities for design and development

Ensure traceability of design and development outputs to design and development

inputs

Availability of resources such technical person, machinery, testing equipment, raw

material etc.

Design shall prepare and develop the product design and also identified the required

resource for this task.

Purchase shall search or identified the vendor or supplier of raw material,

packaging material, any Component etc at best price with best quality.

Marketing shall search the market scope and also monitor the post marketing

surveillance.

MASTER COPY

PASSIM LIFESCIENCES LIMITED

Haryana- ______

DESIGN HISTORY FILE

Name of Product

Piston Nebulizer

Effective Date

27-09-2024

Document Number

PLL/DHF/001

Revision Number

R-01

Next Review Date

Page Number

26-09-2029

Page 3 of 9

•

•

Production shall identified the resource for manufacturing such as machinery,

equipment, and skill man power for production of proto type or samples.

Quality shall verify the produce samples and validate the manufacturing process.

Maintenance shall provide the maintenance support.

Responsibility

Person

1. Mr. Manvinder Singh

2. Mr. Manvinder Singh

Product Design

Purchase

3. Mr. Akhil Mahajan

4. Mr. Akhil Mahajan

5. Mr. Akshit Jaswal

lot

Sign &Date

35371091/2024

Maintenance support

Production of Proto- __________

Quality Control (verification

and validation of Proto-

type/Test lot)

27/09/2024

AH- __________

Aktu

27-09-2014

Jeshi

27-09-2024

2.1 Input from Marketing

2. Design and Development Input

Marketing team has search the availability of product or similar product and found product is

already available worldwide marketing following input is given by marketing manager:

1. Product is available in worldwide market. This is predicate device.

2. Device has good market scope and business opportunity.

3. Device quality should be similar as available in market.

4. Price should be competitive to other reference products available in market.

Name of Person: ____________________

Date & Sign:

2.2 Input from Product Design

(MR)

27/09/2024

1. Product design has identified the material and component of predicate device available in

market. It has various component Tube Mouth piece, Medication Chamber/Cup, Adult Mask,

Child Mask, Air Filter and Material of Construction are Main unit, DC Motor, Base for motor,

Silicone tubes, Power cable, Nebulizer Kit, Anti skid stopper.

2. Need to develop 3-D/2-D design of product outer body in which all component will

MASTER COPY

PASSIM LIFESCIENCES LIMITED

Haryana- ______

DESIGN HISTORY FILE

Name of Product

Piston Nebulizer

Effective Date

27-09-2024

Document Number

PLL/DHF/001

Next Review Date

26-09-2029

Revision Number

R-01

Page Number

Page 4 of 9

assemble.

3. Need to Develop mold.

2.3 Input from Production

Production has given following input:

Material of construction: ______________________________

______________________________

cable, Nebulizer Kit, Anti skid stopper.

2. It is device used to measure the delivery of prescribed medication to adult and infant

population.

3. Source of machinery such as Silicon is required to manufacture product.

4. Skill manpower also required.

Name of Person: ____________________

2.4 Input from QC

Date & Sign:

AKH 709-1024

We have study and review the competitor device and also check the regulatory requirement of

device. We have reach to following conclusion:

Date- __________

,S.NO.- 9 first schedule of Medical Device Rules 2017 which is connected with active

device.

2. Notify device cover under voluntary/mandatory registration till 30.09.2021 after this

cover under compulsory license regime.

3. Device should be accurate and tested.

4. Resource should be available to check device accuracy and functionality.

5

6.

Following instrument should be available for product quality testing such as: Pressure

Gauge flow meter

Device should be tested on functional and physical parameter such as dimension,

appearance, fitment of components.

Refer to Annexure 1 (Test reports)

Name of Person: __________

Jaswal

Designation: ____________

Date &

Closhit

Sign:

3. Design and Development Output

27-09-2014

After review of inputs from all concern departments. Same device of other competitors has been

study and draw the product 2D design. Also identity compatible raw material. Following

material is required to manufacture this device:

1. Main unit should be made of durable plastic material.

2. Tubes should be made of silicon.

MASTER COPY

PASSIM LIFESCIENCES LIMITED

Haryana- ______

DESIGN HISTORY FILE

Name of Product

Piston Nebulizer

Effective Date

27-09-2024

Document Number

PLL/DHF/001

Next Review Date

26-09-2029

Revision Number

R-01

Page Number

Page 5 of 9

Nebulizer Kit- ______________________________

______________________________

Product 2D Design (Piston Nebulizer)

Product Design by:

Mr. Manvinder Singh

Designation:

Director (MR)

4. Design Output Review

Sign & Date:

Wish

27/09/2024

MASTER COPY

PASSIM LIFESCIENCES LIMITED

Haryana- ______

DESIGN HISTORY FILE

Name of Product

Piston Nebulizer

Effective Date

27-09-2024

Document Number

PLL/DHF/001

Revision Number

R-01

Next Review Date

Page Number

26-09-2029

Page 6 of 9

Design Output Review

1) Design output specifies all characteristics of the device that confirm to its safety and

performance for proper use.

2) Design Output cover all functional, Performance, safety and all regulatory requirement.

3) The cover all packaging and labelling requirement

4) Design output meet the input requirement by marketing and Quality that it is similar to

as predicate device

5) Design verification is evident in Process validation batch manufacturing record.

Refer to Annexure- _____________

Design output review by all concern department.

Person: _________________

Responsibility: __________

Date & Sign

AKMY 17-04-2014

5. Design Verification

Verification of Technical Parameter

All technical parameter have been verified. Product has been tested for it functionally and its

performance.

Device characteristics have been verified for:

Parameter

Appearance

Fitment of component

Dimension

Calibration

Specification

Component should be free from burr black spot,

crack.

Should properly fitted

170*135*75 mm

Device should be calibrated

Medication Capacity up to 6ml

Device verified on above technical parameter and complies the all parameters.

We have also verify the regulatory requirement for device verification. This is Class C device

Date- ____________________

Device Rules 2017. We have check the risk management and clinical evaluation applicability.

Risk analysis is carried and control measure are taken to reduce risk and all risk are acceptable

level.

Problem Reporting: ____________________

Calibration report- ______________________________

______________________________

Variable and Pressure Gauge and Flow Meter.

Design Verified By: _________________

Designation: ____________________

Executive

MASTER COPY

Cashi

27-09-2014

PASSIM LIFESCIENCES LIMITED

Plot No. 45, Industrial Area Phase 2, Panchkula, Haryana

DESIGN HISTORY FILE

134113

Name of Product

Piston Nebulizer

Effective Date

27-09-2024

Document Number

PLL/DHF/001

Next Review Date

26-09-2029

Revision Number

R-01

Page Number

Page 7 of 9

6. Design Validation

Device validation of the medical device has been conducted as per the following process flow.

Receiving of Main Unit, motor,

Receiving of

pump, cable wire, air filter)

Covering bag

Receiving of

Nebulizer kit and

Inner box

Assembling of Nebulizer

Calibration & Verification of Nebulizer

Packing of Nebulizer

QC testing

Batch Release

Transfer to FG Store

Three consecutive lot is manufactured of Piston Nebulizer by following the

manufacturing process as mentioned above and test for its physical and functional

performance performance found that produced all batches meet the desired specification each

and every time for the following parameter:

Parameter

Appearance

Fitment of component

Dimension

Calibration

Specification

Component should be free from burr black spot, crack

Should properly fitted

170*135*75 mm

Device should be calibrated Medication Capacity upto 6ml

Its found that produced all batches meet the desired specification each and every time.

Problem Reporting: ______________________________

______________________________

Refer to Annexure- ______________________________

______________________________

slip)

Validation Team

Name of Person

Responsible department

Date & Sign

Production

Neera 2204-2024

Mr. Akshit Jaswal

Quality Control

Mr. Neeraj

MASTER COPY

-09-201

PASSIM LIFESCIENCES LIMITED

Haryana- ______

DESIGN HISTORY FILE

Name of Product

Piston Nebulizer

Effective Date

27-09-2024

Document Number

PLL/DHF/001

Next Review Date

26-09-2029

Revision Number

R-01

Page Number

Page 8 of 9

Mr. Akhil

Maintenance

7. Design and Development Changes

Akki

127-09-2014

6.1 The result of review is pass.

6.2 The result of verification of device is pass.

6.3 The result of validation of device is Pass.

6.4 Device is class C moderate high risk.

6.5 Device performance is good.

6.6 No initiation for design change.

Final Comment: ______________________________

______________________________

by any department. Device design has no changes.

Product Design by: ______________________________

______________________________

Singh

27/09/2024

8. Design Transfer

Design of product and performance of product has been verified and validation. It is pass the all

testing parameter.

1. Design output is record in control format.

2. Product 2-D design is also controlled

3. Device master file as per regulatory requirements has been prepared.

4. Process validation has been also carried.

5. Testing specification and parameter has been controlled in Certificate of Analysis.

6. Production specification, Manufacturing process flow chart, Bill of Material is document

in device master file.

All above said information which is mandatory for device manufacturing is transferred to the

production department in the form of device mater file, design history file, Product 2-D design,

Manufacturing process flow chart, final specification of device is also transfer to production and

quality department.

Design Transfer By:

Mr. Manvinder Singh

Designation:

Director (MR)

Date & Sign:

Received By:

Mr. Akhil

Designation:

Production Manager

MASTER COPY

27/09/2024

Afturi

Date &Şign:

27-09-2014

PASSIM LIFESCIENCES LIMITED

Haryana- ______

DESIGN HISTORY FILE

Name of Product

Piston Nebulizer

Effective Date

27-09-2024

Document Number

PLL/DHF/001

Next Review Date

26-09-2029

Revision Number

R-01

Page Number

Page 9 of 9

Annexure

Annexure

Annexure

Reference

Number

1

23

Test report

Labelling

Batch manufacturing record,

Daily production record,

Material use slip

Attached (Refer to Annexure no. 1).

Attached (Refer to Annexure no. 2)

Attached (Refer to Annexure no. 3)

MASTER COPY

Dr. Odin

PISTON NEBULIZER

100% Pantone 2222C

206x124x175mm

Piston Compressor

Motor

Portable Handheld

Design

Dr. Odin

PISTON NEBULIZER

Strong Working

Air- __________

Low Noise

Content:

Piston nebulizer nebulizer kit (adult mask, child mask,

medication cup tube and mouth piece), ar te non woven

bag and instruction manual

Technical Specification

150 MOUA

1 to 23 PM

52x23x38

Manufactured by:

Passim Lifesciences Ltd.

Phase- __________

Panchkula 134113 (Haryana) - INDIA

For any Complaint/Suggestion

please contact:

Customer Care No.: 1800 309 3009

Timing- ____________________

Email ID: ____________________@__________

Website: _____________

Dr. OdinⓇ

PISTON NEBULIZER

Gray

ONG101

Light Green

Sky Blue

Net Weight:

No. of Pcs.:

11.120 Kg

Gross Weight: __________

Dr. OdinⓇ

PISTON NEBULIZER

THIS SIDE UP

KEEP DRY

8

ONG101

FRAGILE

HANDLE

WITH CARE

USE NO SHARP

INSTRUMENT TO

OPEN

KEEP AWAY

FROM SUNLIGHT

Dr. OdinⓇ

PISTON NEBULIZER

Gray

ONG101

Light Green

Sky Blue

Net Weight:

11.120 Kg

Gross Weight: __________

No. of Pcs.:

8

Dr. OdinⓇ

Piston Nebulizer

Voltage: __________

LOT ONG23001

SN 230100001

MD

i

01-2023

ONG101

ISO

Current: __________

13485:2016

Passim Lifesciences Ltd.

Phase- __________

Panchkula 134113 (HR) - INDIA

Mfg. Lic. No.:

MFG/MD/2022/000633

Email ID: ____________________@__________

Website: _____________

LB/ONG/08-00

190x130mm

Model No.

Invoice No.:

Purchased By: __________

WARRANTY CARD

Lot No.:

Date of Purchase.

Contact No.

Address..

Dealer's Name.

*Terms & Conditions apply

Dealer's Sign & Stamp

Passim Lifesciences Ltd. warrants this units free of any defect in material, workmanship

and operation, under normal use, for a period of one year from the date of purchase. If any

part become defective during the warranty period Passim Lifesciences Ltd. will repair or

replace (if not repairable) the same, free of cost. This shall not apply to any parts that are

considered as expandible or deteriorable in the course of normal use. Passim Lifesciences

Ltd. shall be relieved of any liability and warranty shall cease to apply if:

⚫ This is not used in accordance with the instructions in the operational manual.

It is used with any equipment not complying with the specification of this unit.

It is not regularly maintained.

⚫ The unit is disassembled, repaired or operated by person not authorized by Passim

Lifesciences Ltd.

⚫ Damage caused due to negligency.

⚫ The unit is operated in corrosive materials or in the harmful atmosphere.

The warranty card is not filled completely and produced at the time of warranty claim.

Symbols

MD Medical Device

Refer Instructions Manual

LOT Lot No.

Keep Away From Sunlight

No Trash

Manufactured By

SN Serial No.

Manufacturing Date

Warning/Caution

Manufactured by:

Passim Lifesciences Ltd.

Phase- ____________________

Mfg. Lic. No.: MFG/MD/2022/000633

For any Complaint/ Suggestion please contact:

Customercare Number- ____________________

Email ID: ____________________@__________

IM/ONG/07-00

Dr. Odin

Piston Compresssor

MODEL NO.-ONG101

INSTRUCTION MANUAL

Please read instruction manual before use

190x130mm

Thank you for purchasing the Compressor Nebulizer Model ONG101. It is a

compact medical device designed to deliver the medication prescribed by your

doctor.

Intended Use:

This device is used to deliver medicines to lungs for the treatment of asthma,

allergies and other respiratory disorders. This device is intended for use in a

medical facility, such as hospital, clinic and doctor's office as well as indoors.

Legally certified medical experts, such as doctor, nurse and therapist, or

healthcare personnel or patient under the guidance of qualified medical experts.

This device should be used by patients, who are conscious and can breathe

spontaneously. The user should also be capable of understanding general

machine operation and the content of instruction manual.

Description of device:

This product was developed for the successful treatment of asthma, allergies

and other respiratory disorders. The compressor a stream of air that travels

through the clear tubing to the nebulizer. When the air enters the nebulizer, it

converts the prescribed medication into an aerosol of microscopic droplets

that can easily be inhaled. The Compressor Nebulizer should be used only

as directed. Do not use this device for anything other than as prescribed by

your doctor.

The Compressor Nebulizer Model ONG101 comes with the following

components:

• Compressor (Main Unit)

⚫ Storage Bag

Instruction Manual

Nebulizer Kit (Mouthpiece, Adult Mask (PVC), Child Mask (PVC), Air Tube,

Air Filter

How the nebulizer kit works

The medication that is pumped up through the medication

channel is mixed with compressed air which is generated by a

compressor pump. The compressed air mixed with medication is

turned into fine particles and sprayed when in contact with baffle.

Operating instruction:

1. Place your Compressor Nebulizer on a flat and stable surface. Make sure you

can easily reach the device when you are seated.

2. Make sure the power switch is in the "OFF" state. Plug the device into an

appropriate voltage outlet.

3. Attach one end of the air tube to the air tube connector located on the front of

the outlet point of device.

4. Attach the other end of the air tube base of medicine cup of the nebulizer.

5. Wash your hand. Fill the medicine cup of nebulizer with your medicine. Fill the

medication prescribed by your doctor inside nebulizer. To avoid spills, close the

medicine cup tightly and always hold the mouthpiece straight up and down.

6. Attach the mouthpiece to the medicine cup or mask.

7. Press the power switch ON to start your prescribed treatment.

8. Place the mouthpiece in your mouth. Keep your lips firmly around the

mouthpiece so that all of the medicine goes into your lungs.

9. Breathe through your mouth until all the medicine is used. This takes 10 to 15

minutes. If needed, use a nose clip so that you breathe only through your mouth.

Small children usually do better if they wear a mask.

10. Turn off the machine when done. When treatment is finished, switch off and

unplug the device from the electrical outlet.

11. Wash the medicine cup and mouthpiece with water and air dry until your next

treatment.

Important: ______________________________

______________________________

device before it overheats. When the thermal protector shuts the device off,

please:

a. Switch the device off.

b. Unplug the device from the electrical outlet.

c. Wait 30 minutes for the motor to cool down before another treatment. Make

sure the air openings are not obstructed.

Warnings:

. When treating babies, children or anyone who requires assistance, the

following must be observed without fail. Babies, children or anyone who

requires assistance must be supervised constantly by an adult during inhalation

therapy. This is the only way to ensure safe and effective treatment. In addition,

keep the device out of the reach of those that require supervision such as babies

and children.

• This device is designed only for patients who are conscious and can breathe

spontaneously. This is to the only way to ensure effective treatment and avoid a

choking hazard.

For type, dose and regime of medication, follow the instructions of your

doctor or respiratory therapist.

⚫ Clean and disinfect the nebulizer, mouthpiece and mask before using them for

the first time after purchase, if the device has not been used for a long period of

time, or if more than one person uses the device.

• Always rinse the nebulizer, mouthpiece and mask with clean warm tap water

after disinfecting them.

• Never leave the cleaning solution in the nebulizer or mouthpiece..

• Do not store the air tube while there is moisture or medication remaining inside it.

• Do not use or store the device where it may be exposed to noxious fumes,

flammable gas or volatile substances.

Do not block the air openings. Never place the device where the air openings

may be obstructed during operation. Do not cover the compressor with a blanket

or towel during use.

⚫ Make sure that the nebulizer, mouthpiece and air tube are clean before use.

•Always dispose of any remaining medication after use. Use fresh medication

each time.

Risk of electrical shock

• Never unplug the device with wet hands.

• The compressor and power cord are not waterproof. Do not spill water or

other liquids, on these parts. If liquid does spill on these parts, immediately

unplug the device and wipe off the liquid.

• Do not use or store the device in humid locations such as a bathroom.

• Do not operate the device with a damaged cord or plug.

• Keep the power cord away from heated surfaces.

Caution:

• Make sure that the parts including the filter are attached correctly.

Make sure that the filter is clean. Replace the filter,if the filter has turned

dusty/grey.

Do not tilt the nebulizer kit so the angle of the kit is greater than 45°.

Medication may flow into the mouth. Do not shake the nebulizer kit while using

the device.

• Do not leave the device unattended with babies, children or anyone who

cannot express their consent.

• Do not subject the device or the parts to any strong shocks such as dropping

the device on the floor.

• Do not disassemble or attempt to repair the compressor or power cord.

General Safety Precautions:

Inspect the device and parts before each use for normal operation. Particularly,

be sure to check the following:

That the nebulizer or air tube is not damaged.

. That the cone of the nebulizer is not blocked.

⚫ That the compressor operates normally. The device generates some noise

and vibration caused by the compressor pump, or some noise caused by the

emission of compressed air from the nebulizer. This is normal and does not

indicate a malfunction.

• Operate the device only as intended. Do not use the device for any other

purpose.

• Do not use the device at temperatures greater than +40°C. Make sure that

the air tube is securely attached to the compressor and nebulizer, and does

not come loose. Twist the air tube plug slightly when inserting it into the

connectors to avoid the tube becoming disconnected during use.

Disinfecting the nebulizer:

After the last treatment of the day, follow the procedure below to disinfect your

Compressor Nebulizer unless otherwise specified by your doctor

1. Gently open the nebulizer.

2. Wash the nebulizer, mouthpiece or mask in one part vinegar with three part

distilled water.Make sure solution is enough to submerge into all part nebulizer

kit. Rinse them thoroughly with solution.

3. Dry them with a clean, soft cloth or let them air dry in a clean environment.

4. Reassemble the nebulizer when completely dry. Place these parts in a dry,

sealed container.

Changing the filter:

1. Do not use cotton or any other material .Do not wash or clean the filter. Use

only filter supplied by distributer. Do not operate without a filter.

2. Change the filter when filter turns grey.

Changing Procedure: ______________________________

______________________________

one. Put the filter cap back

Specifications:

Input voltage

Current

Power

Medication Capacity

Particle Size

MMAD

Sound Level

Average Nebulizer Rate

Compressor Pressure Range

Operating Pressure Range

Liter Flow Range

Storage Conditions

Weight

Shelf- ____________________

Warranty

AC 220V, 50Hz

0.5A

45 watt

8ml

0.5 to 10μm

4μm

Less than 60dB at 1 meter

Min. 0.2ml/min.

32 Psi to 45 Psi

16 to 22 Psi

6-9 L/min

Keep it in a cool and dry place

1.4Kg

5 years

1 Year

PASSIM LIFESCIENCES LIMITED

Haryana- ______

Certification of Analysis

Document Title Certification of Analysis for RM/PM/IP/FP

Revision No.

R-02

Effective Date

28-11-2022

Next Review Date

27-11-2024

Page No.

Page 1 of 2

Name of the

Sample:

Piston Nebulizer

ONGIO

STS Number: ______________

Date of

04-02-2023

STP Number: ______________

Sampling:

Sampling

Sopes.

Analytical

Quantity:

Report (AR)

FP- _____

No.:

Material Code:

Lot No.:

ONG101

ONG 23002

Date of

Analysis:

Start Date:

End Date:

04-02-2023 04-02-202

1.0 Test Report of RM/PM/IP/FP:

S. No.

Parameter

Observed

Compliance

Done by

Reading

Yes

No

(Sign &

(✓)

(✓)

Date)

1.

frequency (Hz)

50

2.

Voltage (V)

230

3. flow rate (4/min)

0.4

4. Pressure Range (PST)

45

5.

Product

Check Carton

Justruction

Format- _____________

MASTER COPY

manual,

Serial No.

Sign (QA)

Date:

ISSUED BY QA

C-173

11-01-2023

Biborayo

04-02-20

Bikesan

24-02-20

Biksany

24502-2

Bilray

04,02-20

Bilerany

845027

PASSIM LIFESCIENCES LIMITED

Haryana- ______

Certification of Analysis

Document Title Certification of Analysis for RM/PM/IP/FP

Revision No.

R-02

Effective Date

28-11-2022

2.0 Remark:

Rass

Approved by:

(Sign & Date)

Format- _____________

4/2/23

Next Review Date

Page No.

Page 2 of 2

27-11-2024

Printing, Packs

ging material ✓

NA-

MASTER COPY

Serial No.

Sign (QA)

Date:

ISSUED BY QA

C-173

11-01-2023

Bileray

54502-2012

Format No.

Model

Haryana- ______

ONW102

Lot no. NW22002

Next Review

23-06-2024

Date

F/060/016-00

Effective Date

24-06-2022

DESIGN VALIDATION RECORD

Test report for Piston Nebulizer

Lot size

Date

05

Parameter to be measured

28-06-2022

Sr. No.

Product Sr.

Pressure (psi); >30psi

Flow rate(Litre/min;>6)

Remarks

Tested By

No.

(Pass/Fail/NA)

01

220600001

32 psi

8 l/min

Pass

Auril

02

2206 00002

34 Psi

Slimen

Pass

Alenil

03

856

220600003

36 Psi

81/mon

Pass

Alexis

04

220600004

38 Psi

841302

pass

Alehi

05

220600005

40 Psi

Pass

Are

Prepared by

07

Caif Of

MASTER COPY

Page no 1 of 1

Abhil

Checked by

Serial No.:

Sign (QA)

Date:

ISSUED BY QA

GA- _____

Looney

25-06-2022

Format No.

PASSIM LIFESCIENCES LIMITED

Haryana- ______

Model

ONW102

Lot no. ON W22002

F/060/016-00

Effective Date

24-06-2022

DESIGN VALIDATION RECORD

Next Review

23-06-2024

Date

Test report for Piston Nebulizer

Lot size

Date

05

28-06-2022

Parameter to be measured

Sr. No.

Product Sr.

Pressure (psi); >30psi Flow rate(Litre/min;>6)

Remarks

Tested By

No.

(Pass/Fail/NA)

01

220600006

34

PSS

81/min

Pacs

Akhil

02

220600007

32

Psi

821min

Pass

Akhi

03

220600008

38

Psi

81/min

Pass

Akhil

04

220600009

36

Ps:

811mm

Pass

Aichil

05 220600010

40

Pass

Akhil

Psi

Prepared by

MASTER COPY

Dart Of

Alchil

Checked by

Page no 1 of 1

Serial No.

Sign (QA)

Date:

ISSUED BY QA

GA- _____

25-06-2022

Format No.

PASSIM LIFESCIENCES LIMITED

Haryana- ______

Model

ON WI02

Lot no. ONW 22003

F/060/016-00

Effective Date

24-06-2022

DESIGN VALIDATION RECORD

Test report for Piston Nebulizer

Lot size

Date

Parameter to be measured

Next Review

23-06-2024

Date

Sr. No.

Product Sr.

Pressure (psi); >30psi Flow rate(Litre/min;>6)

No.

Remarks

(Pass/Fail/NA)

Tested By

01

220600011

36 PS

81175

Pass

Akail

02

220600012

32

Psi

81/mm

Paes

Achil

03

220600013

34 PSF

Pass

Akhi

04

220600014

40 Rs

11/mm

pacs

Alch

05 220600015

38 PSE

81/mi

Pass

Architi

Prepared by

MASTER COPY

Cary Of

Page no 1 of 1

Alchil

Checked by

Serial No.:

Sign (QA)

Date:

ISSUED BY QA

A003

t

25-06-2022

Format No.

PASSIM LIFESCIENCES LIMITED

Haryana- ______

Model:

Lot no.:

0NW102

ONW 22001

Parameter to be checked

F/060/08-00

Effective Date

24-06-2022

DESIGN VERIFICATION RECORD

Final Visual Inspection Report for Piston Nebulizer

Lot size:

05

Date:

Next Review

23-06-2024

Date

28-06-2022

Product Sr.

Check

Check

no.

for

switch

Scratch/

on/off

Check

filter

cap

Check

feet

Check

all

Check

Check

Check

Remarks

Product

medication

user

Checked

By

screw

label

kit all part

manual

dent/

body

fitting

✓

220600001

220600002

2206 00003

220600004

220600005

Prepared by

broken

✓

✓

✓

Casy ext

55555

L

☑

Page no 1 of 1

MASTER COPY

2

✓

L

OK

Akhil

OK Akhil

OK

OK

Alen

Alen

OK Aren

Akhil

Checked by

Akhil

Serial No.:

Sign (QA)

Date:

ISSUED BY QA

A-001

Lorey

25-06-2022

Format No.

PASSIM LIFESCIENCES LIMITED

Haryana- ______

F/060/08-00

Effective Date

24-06-2022

DESIGN VERIFICATION RECORD

Model:

Lot no.:

ONW102

ONW 22002

Parameter to be checked

Final Visual Inspection Report for Piston Nebulizer

Lot size:

05

Next Review

23-06-2024

Date

Date- ___/___/_______

Check

Remarks

user

Checked

By

manual

Product Sr.

Check

Check

no.

for

switch

Scratch/

on/off

Check

filter

cap

Check

feet

Check

all

Check

Product

screw

label

dent/

Check

medication

kit all part

body

broken

fitting

22060000 6

220600007

✓

✓

✓

✓

220600008

220600009

£206.00010

✓

D

✓

✓

L

✓

✓

✓

✓

✓

Prepared by

Dart Dy

Page no 1 of 1

MASTER COPY

OK Akhil

OKA

OK Akhil

OK Aen

OK How

Arh

Checked by

Archit

Serial No.

Sign (QA)

Date:

ISSUED BY QA

2A-002

25-06-2022

Format No.

PASSIM LIFESCIENCES LIMITED

Haryana- ______

F/060/08-00

Effective Date

24-06-2022

Model:

Lot no.:

0NW102

ONW 22003

Parameter to be checked

DESIGN VERIFICATION RECORD

Final Visual Inspection Report for Piston Nebulizer

Lot size: _____

Date:

Product Sr.

Check

Check

no.

for

switch

Scratch/

on/off

Check

filter

cap

Check

feet

Check

all

Check

Product

screw

label

dent/

broken

body

fitting

220600011

220600012

220600013

220600014

220600015

Prepared by

Don't Ort

✓

✓

Page no 1 of 1

MASTER COPY

Next Review

23-06-2024

Date

28-06-2022

Check

Remarks Checked

By

Check

medication

kit all part

user

manual

SS SSS

OK

Alent

OK

Alehil

ok

Ales

OK

Allis

OK Arch's

Checked by

Alzhis.

20.06.2022

Serial No.

Sign (QA):

ISSUED BY QA

QA- _____

Son

Batch Manufacturing Record- __________

Format No. PLL/BMR/01-03

Effective Date

13-06-2024 Next Review Date 12-06-2029

BMR Document No.:

MFR No.:

Product Generic Name:

Model No.:

Lot/Batch No.:

Lot Size:

PLL- __________

PLYMFRI 006

Piston Nebulizer

ONS 105

ONS 105 24009

2070 pcs.

Product Serial No.:

Mfg. Date:

Storage condition:

Date of Batch Initiation:

Date of Batch Completion:

Shelf life:

24070900001

24070902070

To

July- _____

Keep

it in a cool and dry placs.

03-08-2024

07-08-2024

-

NA

-

Page 1 of 7

MASTER COPY

Serial No.

ISSUED BY QA

PR- _____

Sign (QA)

Date:

13-2024

Batch Manufacturing Record- __________

Format No. PLL/BMR/01-03

Effective Date 13-06-2024

Next Review Date 12-06-2029

1.0 Bill Of Raw Material:

Material Requisition Slip (F/043/03-current revision) is attached.

Material Code

Description

Standard Quantity Total Quantity

per product

as per Lot

Rm595

Power Cord

2070x1

Rm 182

Rm 183

RM 597

RM185

Back Stopper Rubber

Air filter round shape

Luutch on/off

4

207074

1

2070 XI

1

2070 XI

Silicon Tube

400 mts

400 mtr

RM18

RM 187

Cable tie

Ss. Spring

2

2070x2

3

RM 188

Sonewsph Type sk

6x13

4

2070 X3

2070x4

RM 189

Lubber pad

4

2070x4

RM 190

Medication kit

1

2070X1

RM 191

Compressor Motor

2070x1

RM312

Top Cabinet

2070x1

RM603

Bottom Cabinet

1

2070 XI

RM606

Air Filter Cap blue

2070XI

Rm 195

Rubber band

·500kg

0.500kg

RM 594

Glue stick

6 plat

6Pkt

RM 370

Insulation Tape

12 PC

12pc

NA

Page 2 of 7

MASTER COPY

Serial No.

ISSUED BY QA

PR- _____

Sign (QA)

Date:

13-06-2024

Batch Manufacturing Record- __________

Format No. PLL/BMR/01-03

Effective Date 13-06-2024 Next Review Date

12-06-2029

2.0 Bill Of Packing Material:

Material Requisition Slip (F/043/03-current revision) is attached.

Standard Quantity

Material Code

Total Quantity

Description

per product

PM 496

Instruction Manual

1

Pm124

Pm125

Pm 335

pm224

pm 233

PM247

Polybag 5x7

Non wouch bag

Carton

Comugated box

Surface Cream

Air Filter round shape

(Accessories)

6 kg

as per Lot

2070X1

61g

2070 XI

2070X1

2070-12-173

173 boxes

10

Pcs

1

2070 XI

3.0 Assembly Instructions:

NA

The quality control person shall provide the line clearance by checking the below parameter

to initiate the batch manufacturing:

in- ____________

packed finished product have been removed from the line/area.

3.2) Verify & ensure that assembly line/ production area are properly cleaned.

Page 3 of 7

MASTER COPY

L

Serial No.

Sign (QA)

Date:

ISSUED BY QA

PR- _____

13-06-2024

Batch Manufacturing Record- __________

Format No. PL/BMR/01 0. Effective Date

2024 Next Review Date

12.06 2929

Line Clearance For Manufacturing Process

Manufacturing/Packing area cleaned and free from materials, containers of previous product

(Yes/No):

Previous Product

NA

Area cleaned as per SOP No. PLL/SOP/041

Checked By / Date:

(Production)

0308

Batch/Lot No

Done By /Date

NA

THE

Certified By /Date:

Bitter

(QC)

03-084024

4.0 Assembly:

Assembly Start

S. No.

Date

Time

Date

Assembly End

Time

No. of Product

Done Checked

Assembled

by

by

1

03-08-2024 09:30

03-08-2024 18:30

517 PCs.

Indeep Achi

2 05-08-2024 09:30

3

05-08-2024 18:30

06-08-2024 09:30 06-08-2024 18:30

517 pus

smdlip Akki

518 p

4

07-18-2024 09:30 07-08-2024 18:30

518 ICS

Sondery Desi

Sundeep Aktu

5

6

7

8

9

10

11

12

13

14

15

VA

5.0 Packing Instruction:

The quality control person shall provide the line clearance by checking the below parameter

to initiate the packing of the assembled product:

MASTER COPY

Page 4 of 7

Serial No.

Sigu (QA)

Date:

ISSUED BY QA

2012

13-01-2024

Batch Manufacturing Record- __________

Format No. PLL/BMR/01-03

Effective Date 13-06-2024

Next Review Date 12 06-2029

5.1) Verify the cleaning status of area and ensure all sort of unwanted packing material have

been removed from the line/area.

5.2) Verify & ensure that packing line is properly cleaned and dust free.

Line Clearance For Packing Process

Manufacturing/Packing area cleaned and free from materials, containers of previous product

(Yes/No):

Previous Product:

NA

Area cleaned as per SOP No. PLL/SOP/041

Date: __________

(Production)

03-08-2024....

6.0 Packing:

Batch/ Lot No:

NA

Done By /Date:

Hidrur

Date: __________

(QC)

Packing start

Packing end

No. of product

Done Checked

S. No.

packed

by

by

Date

Time

Date

Time

1

03-08-2024 09:30 03-08-2024

18:30

517 Pcs fondeep Aktuel

2

05-08-2024 09130

05-08.

8.2024

18:30

517 pcs

Sanders Actu

3

06-08-2024 09:30 06-08-2024

18:30

518 pcs

4

07-08-2024

09:30 07-08-2024

18:30

Sundeep Akku

518 pes' Sendup Actul

5

6

78

9

10

7.0 QC Sample Quantity:

21PC3-

(CoA copy is attached with the BMR)

8.0 Transfer Finished Goods to Finished Goods Store

QC release Status:

foleased...

Transfer Ticket- ______________________________

______________________________

MASTER COPY

Page 5 of 7

Serial No.

ISSUED BY QA

PR- _____

Sign (QA)

Date:

13-06-2024

Batch Manufacturing Record- __________

Format No. PLL/BMR/01-03

Effective Date 13-06-2024 Next Review Date

9.0 Reconcilation Of Material:

S. No. Material

Code

1

2

3

4

5

6

7

8

9

12-06-2029

Description

Remaining

Quantity

Action Taken

Reference

Record

MA/RN

Attached

10

Note: ______________________________

______________________________

any).

100%

10.0 Theoretical Yield:

100%

11.0 Actual Yield:

NIL

12.0 Rejection:

13.0 Control Sample:

02 pcs.

14.0 Date of Completion:

07-08-2024

Final test report number and Date- ____________

Page 6 of 7

MASTER COPY

Serial No.

Sign (QA)

Date:

ISSUED BY QA

PR012

13-06-2024

Batch Manufacturing Record- __________

Format No. PLL/BMR/01-03

Effective Date 13-06-2024 Next Review Date

12-06-2029

16.0 Attachment of Labels:

(A4 sheet to be attached in case of shortage of space for labels)

Kushal

Kushal

Prepared By

(Production Officer)

MASTER COPY

Dr. Odin

Piston Nebulizer

ONS105

Voltage: __________

LOT ONS10524009

SN 24070900001

07-2024

MDA

Current: __________

ISO

1345

Passim Lifesciences Ltd.

Phase- __________

Panchkula- _________________

Mfg. Lic. No.

MPO/MD/2012/000633

Email ID: ____________________@__________

Website: _____________

Actui

Verified By

(Production Incharge)

Page 7 of 7

kshut

Verified By

(Quality Incharge)

Serial No.

Sign (QA)

Date:

ISSUED BY QA

PR- _____

13-06-2024

Batch Manufacturing Record- __________

Format No. PLL/BMR/01-03

Effective Date 13-06-2024 Next Review Date 12-06-2029

BMR Document No.:

MFR No.:

Product Generic Name:

Model No.:

Lot/Batch No.:

Lot Size:

Product Serial No.:

Mfg. Date:

Storage condition:

Date of Batch Initiation:

Date of Batch Completion:

Shelf life:

PLL BMR ONS 24-075

PLL MFR 006

Piston Nebulizer

ONS 105

6N510524010

2020 pcs.

24071000001 To 24071002020

July- _____

Keep it in a cool and dry place

25-09-2024

27-09-2024

NA-

Page 1 of 7

MASTER COPY

ISSUED BY QA

PR- _____

Sign (QA)

Date- ___/___/_______

Batch Manufacturing Record- __________

Format No. PLL/BMR/01-03 Effective Date 13-06-2024 Next Review Date 12-06-2029

1.0 Bill Of Raw Material:

Material Requisition Slip (F/043/03-current revision) is attached.

Material Code

Description

Standard Quantity Total Quantity

per product

as per Lot

RM595

Power Cord

1

RM182

Back Stopper Rubber

4

RM183

RM 597

Air filter round shope

2020 XI

2020x4

2020 XI

witch on/off

1

2020 XI

RM 185

Silicon Tube

400 mtr

400 mtr

RM186

RM 187

S.s.

Cable fring

2

2020 X2

3

2020X3

RM 188

Screw Ph Jupe 6x13

4

2020X4

RM 189

Rubber fad

4

20.2014

RM 190

Medication kit

2020X1

Rm 191

RM 312

Top Cabinet

AM 603

Compressor Motor

Bottom Cabinet

2020X1

2020X1

2020X1

RM606

Air Filter

Cop

Blue

2020X1

RM 195

Rubber band

0'

• 500kg

0.500kg

RM594

Glue Stick

6 pkt

6 pret

Rm 370

Insulation Tape

10

10PG

Page 2 of 7

MASTER COPY

ISSUED BY QA

Serial No.

PR- _____

Sign (QA)

Date- ___________

Batch Manufacturing Record- __________

Format No.

PLL/BMR/01-03

2.0 Bill Of Packing Material:

Effective Date 13-06-2024 Next Review Date 12-06-2029

Material Requisition Slip (F/043/03 current revision) is attached.

Material Code

PM 496

PM124

PM 125

pm 335

pm224

pm233

Pm247

Description

Instruction Manual

Polybag 5X7

Non woven bag

Carton

Comugated box

Surface Cream

Air filter round shape

Standard Quantity Total Quantity

per product

1

1

1

kg

2990+12

10 PC

as per Lot

2020X1

6kg

2020X1

2020X1

169 pes

Lo pe

2020 X1

Accessories

NA

3.0 Assembly Instructions:

The quality control person shall provide the line clearance by checking the below parameter

to initiate the batch manufacturing:

in- ____________

packed finished product have been removed from the line/area.

3.2) Verify & ensure that assembly line/ production area are properly cleaned.

Page 3 of 7

MASTER COPY

ISSUED BY QA

Serial No.

PB- _____

Sign (QA):

Date- ___/___/_______

Batch Manufacturing Record- __________

Format No. PLL/BMR/01-03 Effective Date 13-06-2024 Next Review Date

12-06-2029

Line Clearance For Manufacturing Process

Manufacturing/Packing area cleaned and free from materials, containers of previous product

(Yes/No):

Previous Product:

NA

Checked By Date:

Area cleaned as per SOP No. PLL/SOP/041

(Production)

Nieve

25-09-2024

Batch/ Lot No:

NA

Done By /Date:

Date: __________

(QC)

25042047..

Susat

4.0 Assembly:

Assembly Start

Assembly End

No. of Product

Done

S. No.

Assembled

by

Checked

by

Date

Time

Date

Time

1

25-09-2014 09:30 25-09-2024 18:30

674 pcs

2

3

26-09-2024 09:30 26-09-224 18:30

27-09-2014 09:30 27-09-2024

674 Pcs

18:30 672 pcs. Vikas

Vikas Aktu

Vikas Akhil

Vikas Aptu

4

10

5

6

IN

7

8

9

0123

11

NA

15

14

15

5.0 Packing Instruction:

The quality control person shall provide the line clearance by checking the below parameter

to initiate the packing of the assembled product:

MASTER COPY

Page 4 of 7

ISSUED BY QA

Serial No. R-028

Sign (QA)

01-09-2024

Batch Manufacturing Record- __________

Format No. PLL/BMR/01-03 Effective Date 13-06-2024 Next Review Date

12-06-2029

5.1) Verify the cleaning status of area and ensure all sort of unwanted packing material have

been removed from the line/area.

5.2) Verify & ensure that packing line is properly cleaned and dust free.

Line Clearance For Packing Process

Manufacturing/Packing area cleaned and free from materials, containers of previous product

(Yes/No):

Previous Product:

NA

Checked By Date: __________

Area cleaned as per SOP No. PLL/SOP/041

(Production)

Nevio

25092024

Batch/ Lot No:

NA

Done By /Date:

Certified By /Date:

(QC)

tcy

Date: __________

25092024

6.0 Packing:

Packing start

Packing end

S. No.

No. of product

packed

Done Checked

by

by

Date

Time

Date

Time

1

25.09.2024 09:30 25.09-2024

18130

2

26-09-2024 09:30

26-09-2024

18:30

3

27-09-2024 09:30 27-09-2024

18:30

674 pcs Vikas Adic

674 Pcs Vikas Aktu

672 pcs. Vikal. Alchi

4.

5

67

00

8

9

10

10

·NA

QC Sample Quantity: __________

(CoA copy is attached with the BMR)

8.0 Transfer Finished Goods to Finished Goods Store

QC release Status: __________

Released

Transfer Ticket- ______________________________

______________________________

MASTER COPY

Page 5 of 7

ISSUED BY QA

PR- _____

Sign (QA)

Photo 01-09-2024

Batch Manufacturing Record- __________

Format No. PLL/BMR/01-03

9.0 Reconcilation Of Material:

S. No.

Material

Code

1

23

Effective Date

13-06-2024 Next Review Date

12-06-2029

Description

Remaining

Action Taken

Quantity

Reference

Record

4.

567

8

9

10

MRN

Attached

Note: ______________________________

______________________________

any).

10.0 Theoretical Yield:

30.

Actual Yield: __________

100%

12.0 Rejection:

NIL

13.0 Control Sample:

02

Pes

27-09-2024

14.0 Date of Completion:

15.0 AR No. /Final test report number and Date:

Page 6 of 7

MASTER COPY

24/PP/24-0106

ISSUED BY QA

PR- _____

Sign (QA)

Date- ___/___/_______

Batch Manufacturing Record- __________

Effective Date 13-06-2024

Next Review Date

12-06-2029

Format No.

PLL/BMR/01-03

16.0 Attachment of Labels:

(A4 sheet to be attached in case of shortage of space for labels)

Kishal

Dr. Odin

Piston Nebulizer

ONS105

Voltage: __________

LOT ONS10524010

SN 24071000001

07-2024

MDA

Current: __________

45

Passim Lifesciences Ltd.

Phase- __________

Panchkula- __________________

My Lie. Na

MFG/MD/2022/000433

Email ID: ____________________@__________

Website: _____________

Nevroy

Prepared By

(Production Officer)

Verified By

(Production Incharge)

MASTER COPY

Page 7 of 7

Verified By

(Quality Incharge)

ISSUED BY QA

PR- _____

Sign (04)

Date- ___/___/_______

Batch Manufacturing Record- __________

Format No. PLL/BMR/01-03

Effective Date

13-06-2024 Next Review Date

12:06-2029

BMR Document No.:

MFR No.:

Product Generic Name:

Model No.:

Lot/Batch No.:

Lot Size:

Product Serial No.:

Mfg. Date:

Storage condition:

Date of Batch Initiation:

Date of Batch Completion:

Shelf life:

PLL- ______________

PLL MFR 006

Piston Nebulizer.

ONS105

ONS 10524011

2030 pcs.

24071100001

24071102030

To

July- _____

keep it in a cool and dry place.

07-09-2024

09-09-2024

NA-

Page 1 of 7

MASTER COPY

ISSUED BY QA

Bedal No

Sign (QA)

Date:

PR- _____

13-06-2024

Batch Manufacturing Record- __________

Format No. PLL/BMR/01-03

1.0 Bill Of Raw Material:

Effective Date 13-06-2024 Next Review Date

12-06-2029

Material Requisition Slip (F/043/03-current revision) is attached.

Standard Quantity Total Quantity

Material Code

Description

RM 595

Power Cord

RM 182

Bm 183

RM 591

RM 185

RM 186

Back Stopper Rubber.

An Filter round shape

on/off switch

Silicon Tube

S.S. Specing

per product

1

4

as per Lot

2030X1

2030 X4

2030 X1

1

2030X1

400 Mtr

400 mt

2

2030X2

Rm187

Cable Te

3

2030X3

RM188

Screw PhType 6x13

4

2030 X4-

RM 189

Rubber Pad

4

203084

RM 190

Medication kit

2030x1

Rm191

Compressor Motor

2030x1

RM312

Top Cabinet

2030x1

Rm602

Bottom Cabomit

1

RM605

Amr Filter Cap

2030x1

2030X1

Rm 195

Rubber band

Rm594

Glue stick

0.500 kg

6PKJ

0.500kg

6 Pkt

RM370

Insulation Tape

10 Pe

10 pc

NA

Page 2 of 7

MASTER COPY

Serial Net

Sign (QA)

Date:

19SUED BY QA

PR- _____

13-06-2024

Batch Manufacturing Record- __________

Format No. PLL/BMR/01-03

Effective Date 13-06-2024 Next Review Date 12-06-2029

2.0 Bill Of Packing Material:

Material Requisition Slip (F/043/03-current revision) is attached.

Material Code

Description

Standard Quantity

per product

Total Quantity

as per Lot

Pm498

Instruction Manual

2030X1

pm124

poly bag 5x7

6 kg

6kg

M125

No won

Pm414

pm224

PM 223

PM247

neoven bag

Carton

Comugated box

Surface Cream

Air filter round shape.

Accessory.

203012

10 PC

2830X1

2030X1

170 boxes

10pc

2030X)

NA

3.0 Assembly Instructions:

The quality control person shall provide the line clearance by checking the below parameter

to initiate the batch manufacturing:

in- ____________

packed finished product have been removed from the line/area.

3.2) Verify & ensure that assembly line/ production area are properly cleaned.

Page 3 of 7

MASTER COPY

Serial Not

Sign (0)

Date:

ISSE FILIY QA

PR- _____

F

13-06-2024

Batch Manufacturing Record- __________

Format No. PLL/BMR/01-03

Effective Date 13-06-2024 Next Review Date

Line Clearance For Manufacturing Process

12-06-2029

Manufacturing/Packing area cleaned and free from materials, containers of previous product

(Yes/No):

Previous Product:

NA

Area cleaned as per SOP No. PLL/SOP/041

Lot No: __________

Done By /Date:

Date: __________

.....

(Production)

07-09 2014.....

NA

Date: __________

(QC)

07-09-2014".

4.0 Assembly:

Assembly Start

Assembly End

No. of Product

Done

Checked

S. No.

Assembled

by

by

Date

Time

Date

Time

1

07-09-2024 09:30

2

09.09.2024

09:30

07-09-2024 18:30

09-09-2024 18:30

1015 pcs. Curpreet Aut

1015 pcs

Current Delle

3

4

5

67

8 9

10

11

12

13

14

15

5.0 Packing Instruction:

The quality control person shall provide the line clearance by checking the below parameter

to initiate the packing of the assembled product:

MASTER COPY

Page 4 of 7

Serial No

Sign (QA

Date:

ISSUED BY QA

023

f

13-06-2024

Batch Manufacturing Record- __________

Format No. PLL/BMR/01-03

Effective Date 13-06-2024 Next Review Date 12-06-2029

5.1) Verify the cleaning status of area and ensure all sort of unwanted packing material have

been removed from the line/area.

5.2) Verify & ensure that packing line is properly cleaned and dust free.

Line Clearance For Packing Process

Manufacturing/Packing area cleaned and free from materials, containers of previous product

(Yes/No):

Previous Product:

NA

Area cleaned as per SOP No. PLL/SOP/041

Date: __________

(Production)

.. Kusshay

17.09.2024.

NA

Batch/Lot No:

Done By /Date:

Date: __________

(QC)

27-09-2024

6.0 Packing:

Packing start

Packing end

S. No.

No. of product

packed

Done Checked

by

by

Date

Time

1

07-092024 09130

Date

07-09-2024

Time

18:30

2

09-09-2024 09:30 09-09-2024 18:30

1015 Pcs.

Current Achi

1015 pcs. Current Acrui

3

4

5

6

7

8

9

10

QC Sample Quantity: __________________

(CoA copy is attached with the BMR)

8.0 Transfer Finished Goods to Finished Goods Store

... Released....

QC release Status: __________

Transfer Ticket- ______________________________

______________________________

MASTER COPY

Page 5 of 7

Serial Nost

Sign

Date:

ESSED BY O

PR- _____

13-06-2024

Batch Manufacturing Record- __________

Format No. PLL/BMR/01-03 Effective Date

13-06-2024

9.0 Reconcilation Of Material:

S. No.

Material

Code

1

2

3

4

5

6

7

8

9

Next Review Date 12-06-2029

Description

Remaining

Action Taken

Quantity

Reference

Record

NO

10

Note: ______________________________

______________________________

any).

100%

10.0 Theoretical Yield:

11.0 Actual Yield:

100%

NIL

12.0 Rejection:

13.0 Control Sample:

-09-09-2024 02 pcs.

14.0 Date of Completion:

09-09-2024

15.0 AR No. /Final test report number and Date:

Page 6 of 7

MASTER COPY

QC/FP/24-0101

Serial

Sign (0)

Dat

PR- _____

13-06-2024

Batch Manufacturing Record- __________

Format No. PLL/BMR/01-03 Effective Date

13-06-2024 Next Review Date 12-06-2029

16.0 Attachment of Labels:

(A4 sheet to be attached in case of shortage of space for labels)

Kesha

Dr. Odin

Piston Nebulizer

ONS105

Voltage: ____________________

LOT ONS10524011

SN 24071100001

07-2024

MDA

Passim Lifesciences Ltd.

Phase- __________

Panchkula- _________________

Mly. Lic. No.:

MFG/MD/2011/000633

Email ID: ____________________@__________

Website: _____________

Kehal

Prepared By

(Production Officer)

MASTER COPY

Aemi

Verified By

(Production Incharge)

Page 7 of 7

afshet

Verified By

(Quality Incharge)

Serial

Sign (0

Date:

LMED BY OA

Ph023

f

13-06-2024

## Instructions:

- Fill in all fields marked with underscores

- Replace _____ with appropriate information

- Ensure all required fields are completed