# 🧠 INTELLIGENT BLANK TEMPLATE (FROM PDF)

📄 Original PDF: DHF..pdf

🎯 Intelligent Analysis Confidence: 100.0%

📊 Detected: 5 tables, 1399 headers, 305 fields

📝 Original layout and structure recreated. Fill in all blank fields marked with underscores.

================================================================================

Customer Care No.: _______________ _______________

(Timing: _______________

Email ID: ________@_______.___

Customercare Number: __________________________________________

| PLL | BMR | ONS | 24-060 |
| --- | --- | --- | --- |

| PLL | BMR/ONS 24-071 |
| --- | --- |

___/___/_____ 18: ________

# PASSIM LIFESCIENCES LIMITED

## Plot No. 45, Industrial Area Phase 2, Panchkula, Haryana - 134113

# DESIGN HISTORY FILE

Name of Product

## Piston Nebulizer

## Effective Date

___/___/_____

## Document Number

# PLL/DHF/001

## Next Review Date

___/___/_____

## Revision Number

# R-01

## Page Number

Page 1 of 9

## Design History File

## (Piston Nebulizer)

## Manufacturing Site:

## Passim Lifesciences Limited

## Plot No. 45, Industrial Area Phase 2,

## Panchkula, Haryana - 134113

This design history file is prepared as per requirement of ISO _____: ________, clause 7.3 design and

development requirement. It is maintained at above said manufacturing site.

## Prepared By:

## Fira

___/___/_____

# MASTER COPY

# MAST

## Approved by:

___/___/_____

# PASSIM LIFESCIENCES LIMITED

## Plot No. 45, Industrial Area Phase 2, Panchkula, Haryana - 134113

# DESIGN HISTORY FILE

Name of Product

## Piston Nebulizer

## Effective Date

___/___/_____

## Document Number

# PLL/DHF/001

## Next Review Date

___/___/_____

## Revision Number

# R-01

## Page Number

Page 2 of 9

# 1. Design and Development Planning

Description of Project: ______________________________

planned due to emerging need of active medical device and its accessories in market to fulfil

the need of user at affordable prices and for better business growth.

Design planning of Piston Nebulizer was initiated to add new product range also and

upcoming need to active medical device accessories.

Requirement was raised by the marketing team to see the product market growth.

## Design and Development Planning Approach:

Following design and development stages applied to the product Piston Nebulizer :

## Design & Development Planning +

Control of Design & Development Changes

## Design & Development Inputs

## Design & Development Output

## Design & Development Review

## Design & Development

## Verification

## Design & Development

## Validation

•

•

The review(s) needed at each design and development stage

The verification, validation, and design transfer activities that are appropriate at each

design and development stage.

Define the responsibilities and authorities for design and development

Ensure traceability of design and development outputs to design and development

inputs

Availability of resources such technical person, machinery, testing equipment, raw

material etc.

Design shall prepare and develop the product design and also identified the required

resource for this task.

Purchase shall search or identified the vendor or supplier of raw material,

packaging material, any Component etc at best price with best quality.

Marketing shall search the market scope and also monitor the post marketing

surveillance.

# MASTER COPY

# PASSIM LIFESCIENCES LIMITED

## Plot No. 45, Industrial Area Phase 2, Panchkula, Haryana - 134113

# DESIGN HISTORY FILE

Name of Product

## Piston Nebulizer

## Effective Date

___/___/_____

## Document Number

# PLL/DHF/001

## Revision Number

# R-01

## Next Review Date

## Page Number

___/___/_____

Page 3 of 9

•

•

Production shall identified the resource for manufacturing such as machinery,

equipment, and skill man power for production of proto type or samples.

Quality shall verify the produce samples and validate the manufacturing process.

Maintenance shall provide the maintenance support.

## Responsibility

## Person

# 1. Mr. Manvinder Singh

# 2. Mr. Manvinder Singh

## Product Design

## Purchase

# 3. Mr. Akhil Mahajan

# 4. Mr. Akhil Mahajan

# 5. Mr. Akshit Jaswal

lot

## Sign &Date

________/____

Maintenance support

Production of Proto-type/Test

Quality Control (verification

and validation of Proto-

type/Test lot)

___/___/_____

# AH-09-2014

## Aktu

___/___/_____

## Jeshi

___/___/_____

2.1 Input from Marketing

# 2. Design and Development Input

Marketing team has search the availability of product or similar product and found product is

already available worldwide marketing following input is given by marketing manager:

# 1. Product is available in worldwide market. This is predicate device.

# 2. Device has good market scope and business opportunity.

# 3. Device quality should be similar as available in market.

# 4. Price should be competitive to other reference products available in market.

Name of Person: _________________________________________

## Date & Sign:

2.2 Input from Product Design

# (MR)

___/___/_____

1. Product design has identified the material and component of predicate device available in

market. It has various component Tube Mouth piece, Medication Chamber/Cup, Adult Mask,

Child Mask, Air Filter and Material of Construction are Main unit, DC Motor, Base for motor,

Silicone tubes, Power cable, Nebulizer Kit, Anti skid stopper.

2. Need to develop 3-D/2-D design of product outer body in which all component will

# MASTER COPY

# PASSIM LIFESCIENCES LIMITED

## Plot No. 45, Industrial Area Phase 2, Panchkula, Haryana - 134113

# DESIGN HISTORY FILE

Name of Product

## Piston Nebulizer

## Effective Date

___/___/_____

## Document Number

# PLL/DHF/001

## Next Review Date

___/___/_____

## Revision Number

# R-01

## Page Number

Page 4 of 9

assemble.

# 3. Need to Develop mold.

2.3 Input from Production

## Production has given following input:

1. Material of construction: ______________________________

cable, Nebulizer Kit, Anti skid stopper.

2. It is device used to measure the delivery of prescribed medication to adult and infant

population.

# 3. Source of machinery such as Silicon is required to manufacture product.

# 4. Skill manpower also required.

Name of Person: _________________________________________________

2.4 Input from QC

## Date & Sign:

# AKH 709-1024

We have study and review the competitor device and also check the regulatory requirement of

device. We have reach to following conclusion: _______________

1. This is Class C low risk device File No. 29/Misc/3/____-DC (292), Date: _______________ _______________ _______________ _______________ _______________ _______________ _______________ _______________ _______________ _______________ _______________ _______________ _______________ _______________ _______________ _______________ _______________ _______________ _______________ _______________ _______________ _______________ 15.05.201

,S.NO.- 9 first schedule of Medical Device Rules ____ which is connected with active

device.

2. Notify device cover under voluntary/mandatory registration till 30.09.____ after this

cover under compulsory license regime.

# 3. Device should be accurate and tested.

# 4. Resource should be available to check device accuracy and functionality.

5

6.

Following instrument should be available for product quality testing such as: Pressure

Gauge flow meter

Device should be tested on functional and physical parameter such as dimension,

appearance, fitment of components.

Refer to Annexure 1 (Test reports)

Name of Person: _______________

## Jaswal

Designation: _______________

## Date &

## Closhit

## Sign:

# 3. Design and Development Output

___/___/_____

## After review of inputs from all concern departments. Same device of other competitors has been

study and draw the product 2D design. Also identity compatible raw material. Following

material is required to manufacture this device: _______________

# 1. Main unit should be made of durable plastic material.

# 2. Tubes should be made of silicon.

# MASTER COPY

# PASSIM LIFESCIENCES LIMITED

## Plot No. 45, Industrial Area Phase 2, Panchkula, Haryana - 134113

# DESIGN HISTORY FILE

Name of Product

## Piston Nebulizer

## Effective Date

___/___/_____

## Document Number

# PLL/DHF/001

## Next Review Date

___/___/_____

## Revision Number

# R-01

## Page Number

Page 5 of 9

# 3. Nebulizer Kit - Medication cup, mouthpiece, tube, adult and child mask..

## Product 2D Design (Piston Nebulizer)

## Product Design by:

## Mr. Manvinder Singh

## Designation:

Director (MR)

# 4. Design Output Review

## Sign & Date:

## Wish

___/___/_____

# MASTER COPY

# PASSIM LIFESCIENCES LIMITED

## Plot No. 45, Industrial Area Phase 2, Panchkula, Haryana - 134113

# DESIGN HISTORY FILE

Name of Product

## Piston Nebulizer

## Effective Date

___/___/_____

## Document Number

# PLL/DHF/001

## Revision Number

# R-01

## Next Review Date

## Page Number

___/___/_____

Page 6 of 9

## Design Output Review

1) Design output specifies all characteristics of the device that confirm to its safety and

performance for proper use.

2) Design Output cover all functional, Performance, safety and all regulatory requirement.

3) The cover all packaging and labelling requirement

4) Design output meet the input requirement by marketing and Quality that it is similar to

as predicate device

5) Design verification is evident in Process validation batch manufacturing record.

Refer to Annexure-2 (Labelling)

## Design output review by all concern department.

## Person :Mr. Akhil Mahajan

## Responsibility: Production

## Date & Sign

# AKMY 17-04-2014

# 5. Design Verification

Verification of Technical Parameter

All technical parameter have been verified. Product has been tested for it functionally and its

performance.

## Device characteristics have been verified for:

## Parameter

## Appearance

Fitment of component

## Dimension

## Calibration

## Specification

Component should be free from burr black spot,

crack.

Should properly fitted

170*135*75 mm

Device should be calibrated

Medication Capacity up to 6ml

Device verified on above technical parameter and complies the all parameters.

We have also verify the regulatory requirement for device verification. This is Class C device

device File No. 29/Misc/3/2017-DC (292), Date: ___/___/_____

Device Rules ____. We have check the risk management and clinical evaluation applicability.

Risk analysis is carried and control measure are taken to reduce risk and all risk are acceptable

level.

Problem Reporting: _________________________

Calibration report- Refer to the calibration file for the report, Instruments used-Frequency

Variable and Pressure Gauge and Flow Meter.

## Design Verified By: Mr. Akshit Jaswal

## Designation: Quality Control Date & Sign:

## Executive

# MASTER COPY

## Cashi

___/___/_____

# PASSIM LIFESCIENCES LIMITED

## Plot No. 45, Industrial Area Phase 2, Panchkula, Haryana

# DESIGN HISTORY FILE

______

Name of Product

## Piston Nebulizer

## Effective Date

___/___/_____

## Document Number

# PLL/DHF/001

## Next Review Date

___/___/_____

## Revision Number

# R-01

## Page Number

Page 7 of 9

# 6. Design Validation

Device validation of the medical device has been conducted as per the following process flow.

Receiving of Main Unit, motor,

Receiving of

pump, cable wire, air filter)

Covering bag

Receiving of

Nebulizer kit and

Inner box

Assembling of Nebulizer

Calibration & Verification of Nebulizer

Packing of Nebulizer

QC testing

## Batch Release

Transfer to FG Store

Three consecutive lot is manufactured of Piston Nebulizer by following the

manufacturing process as mentioned above and test for its physical and functional

performance performance found that produced all batches meet the desired specification each

and every time for the following parameter: _______________

## Parameter

## Appearance

Fitment of component

## Dimension

## Calibration

## Specification

Component should be free from burr black spot, crack

Should properly fitted

170*135*75 mm

Device should be calibrated Medication Capacity upto 6ml

Its found that produced all batches meet the desired specification each and every time.

Problem Reporting: ______________________________

Refer to Annexure-3 (Batch manufacturing record,Daily production slip, Material use

slip)

## Validation Team

Name of Person

## Responsible department

## Date & Sign

## Production

## Neera 2204-2024

## Mr. Akshit Jaswal

## Quality Control

## Mr. Neeraj

# MASTER COPY

-09-201

# PASSIM LIFESCIENCES LIMITED

## Plot No. 45, Industrial Area Phase 2, Panchkula, Haryana - 134113

# DESIGN HISTORY FILE

Name of Product

## Piston Nebulizer

## Effective Date

___/___/_____

## Document Number

# PLL/DHF/001

## Next Review Date

___/___/_____

## Revision Number

# R-01

## Page Number

Page 8 of 9

## Mr. Akhil

## Maintenance

# 7. Design and Development Changes

6.1 The result of review is pass.

6.2 The result of verification of device is pass.

6.3 The result of validation of device is Pass.

6.4 Device is class C moderate high risk.

6.5 Device performance is good.

6.6 No initiation for design change.

## Akki

1___/___/_____

Final Comment: ______________________________

## by any department. Device design has no changes.

Product Design by: ______________________________________________________

## Singh

___/___/_____

# 8. Design Transfer

Design of product and performance of product has been verified and validation. It is pass the all

testing parameter.

# 1. Design output is record in control format.

# 2. Product 2-D design is also controlled

# 3. Device master file as per regulatory requirements has been prepared.

# 4. Process validation has been also carried.

5. Testing specification and parameter has been controlled in Certificate of Analysis.

6. Production specification, Manufacturing process flow chart, Bill of Material is document

in device master file.

All above said information which is mandatory for device manufacturing is transferred to the

## production department in the form of device mater file, design history file, Product 2-D design,

Manufacturing process flow chart, final specification of device is also transfer to production and

## quality department.

## Design Transfer By:

## Mr. Manvinder Singh

## Designation:

Director (MR)

## Date & Sign:

## Received By:

## Mr. Akhil

## Designation:

## Production Manager

# MASTER COPY

___/___/_____

## Afturi

## Date &Şign:

___/___/_____

# PASSIM LIFESCIENCES LIMITED

## Plot No. 45, Industrial Area Phase 2, Panchkula, Haryana - 134113

# DESIGN HISTORY FILE

Name of Product

## Piston Nebulizer

## Effective Date

___/___/_____

## Document Number

# PLL/DHF/001

## Next Review Date

___/___/_____

## Revision Number

# R-01

## Page Number

Page 9 of 9

## Annexure

## Annexure

## Annexure

## Reference

## Number

1

23

Test report

## Labelling

Batch manufacturing record,

Daily production record,

Material use slip

Attached (Refer to Annexure no. 1).

Attached (Refer to Annexure no. 2)

Attached (Refer to Annexure no. 3)

# MASTER COPY

## Dr. Odin

# PISTON NEBULIZER

## 100% Pantone 2222C

206x124x175mm

## Piston Compressor

## Motor

## Portable Handheld

## Design

## Dr. Odin

# PISTON NEBULIZER

## Strong Working

## Air-Flow

## Low Noise

## Content:

Piston nebulizer nebulizer kit (adult mask, child mask,

medication cup tube and mouth piece), ar te non woven

bag and instruction manual

## Technical Specification

# 150 MOUA

## 1 to 23 PM

## 52x23x38

## Manufactured by:

## Passim Lifesciences Ltd.

Plot No. 45, Ind. Area, Phase-II,

Panchkula ______ (Haryana) - INDIA

For any Complaint/Suggestion

## please contact:

Website: _______________

Dr. OdinⓇ

# PISTON NEBULIZER

## Gray

# ONG101

## Light Green

## Sky Blue

## Net Weight:

## No. of Pcs.:

## 11.120 Kg

## Gross Weight: 12.000 Kg

Dr. OdinⓇ

# PISTON NEBULIZER

# THIS SIDE UP

# KEEP DRY

8

# ONG101

# FRAGILE

# HANDLE

# WITH CARE

# USE NO SHARP

# INSTRUMENT TO

# OPEN

# KEEP AWAY

# FROM SUNLIGHT

Dr. OdinⓇ

# PISTON NEBULIZER

## Gray

# ONG101

## Light Green

## Sky Blue

## Net Weight:

## 11.120 Kg

## Gross Weight: 12.000 Kg

## No. of Pcs.:

8

## Dr. OdinⓇ

## Piston Nebulizer

## Voltage: 220V

# LOT ONG23001

# SN 230100001

# MD

i

01-____

# ONG101

# ISO

## Current: 0.5A

13485: ________

## Passim Lifesciences Ltd.

Plot No. 45, Ind. Area, Phase - II,

Panchkula ______ (HR) - INDIA

## Mfg. Lic. No.:

# MFG/MD/2022/000633

Email ID: ________@_______.___

Website: _______________

# LB/ONG/08-00

## 190x130mm

## Model No.

## Invoice No.:

## Purchased By:.

# WARRANTY CARD

## Lot No.:

Date of Purchase.

## Contact No.

## Address..

Dealer's Name.

*Terms & Conditions apply

Dealer's Sign & Stamp

Passim Lifesciences Ltd. warrants this units free of any defect in material, workmanship

and operation, under normal use, for a period of one year from the date of purchase. If any

## part become defective during the warranty period Passim Lifesciences Ltd. will repair or

## replace (if not repairable) the same, free of cost. This shall not apply to any parts that are

considered as expandible or deteriorable in the course of normal use. Passim Lifesciences

Ltd. shall be relieved of any liability and warranty shall cease to apply if:

⚫ This is not used in accordance with the instructions in the operational manual.

It is used with any equipment not complying with the specification of this unit.

It is not regularly maintained.

⚫ The unit is disassembled, repaired or operated by person not authorized by Passim

## Lifesciences Ltd.

⚫ Damage caused due to negligency.

⚫ The unit is operated in corrosive materials or in the harmful atmosphere.

The warranty card is not filled completely and produced at the time of warranty claim.

## Symbols

MD Medical Device

## Refer Instructions Manual

LOT Lot No.

## Keep Away From Sunlight

## No Trash

## Manufactured By

SN Serial No.

## Manufacturing Date

## Warning/Caution

## Manufactured by:

## Passim Lifesciences Ltd.

Plot No. 45, Ind. Area, Phase-II, Panchkula -______, Haryana- INDIA

Mfg. Lic. No.: _______________

## For any Complaint/ Suggestion please contact:

Email ID: ________@_______.___

# IM/ONG/07-00

## Dr. Odin

## Piston Compresssor

# MODEL NO.-ONG101

# INSTRUCTION MANUAL

Please read instruction manual before use

## 190x130mm

Thank you for purchasing the Compressor Nebulizer Model ______. It is a

compact medical device designed to deliver the medication prescribed by your

doctor.

## Intended Use:

This device is used to deliver medicines to lungs for the treatment of asthma,

allergies and other respiratory disorders. This device is intended for use in a

medical facility, such as hospital, clinic and doctor's office as well as indoors.

Legally certified medical experts, such as doctor, nurse and therapist, or

healthcare personnel or patient under the guidance of qualified medical experts.

This device should be used by patients, who are conscious and can breathe

spontaneously. The user should also be capable of understanding general

machine operation and the content of instruction manual.

## Description of device:

This product was developed for the successful treatment of asthma, allergies

and other respiratory disorders. The compressor a stream of air that travels

through the clear tubing to the nebulizer. When the air enters the nebulizer, it

converts the prescribed medication into an aerosol of microscopic droplets

that can easily be inhaled. The Compressor Nebulizer should be used only

as directed. Do not use this device for anything other than as prescribed by

your doctor.

The Compressor Nebulizer Model ______ comes with the following

## components:

## • Compressor (Main Unit)

## ⚫ Storage Bag

## Instruction Manual

Nebulizer Kit (Mouthpiece, Adult Mask (PVC), Child Mask (PVC), Air Tube,

## Air Filter

How the nebulizer kit works

The medication that is pumped up through the medication

channel is mixed with compressed air which is generated by a

compressor pump. The compressed air mixed with medication is

## turned into fine particles and sprayed when in contact with baffle.

## Operating instruction:

# 1. Place your Compressor Nebulizer on a flat and stable surface. Make sure you

can easily reach the device when you are seated.

# 2. Make sure the power switch is in the "OFF" state. Plug the device into an

appropriate voltage outlet.

3. Attach one end of the air tube to the air tube connector located on the front of

the outlet point of device.

# 4. Attach the other end of the air tube base of medicine cup of the nebulizer.

5. Wash your hand. Fill the medicine cup of nebulizer with your medicine. Fill the

medication prescribed by your doctor inside nebulizer. To avoid spills, close the

medicine cup tightly and always hold the mouthpiece straight up and down.

# 6. Attach the mouthpiece to the medicine cup or mask.

# 7. Press the power switch ON to start your prescribed treatment.

# 8. Place the mouthpiece in your mouth. Keep your lips firmly around the

mouthpiece so that all of the medicine goes into your lungs.

9. Breathe through your mouth until all the medicine is used. This takes 10 to 15

minutes. If needed, use a nose clip so that you breathe only through your mouth.

Small children usually do better if they wear a mask.

# 10. Turn off the machine when done. When treatment is finished, switch off and

unplug the device from the electrical outlet.

# 11. Wash the medicine cup and mouthpiece with water and air dry until your next

treatment.

Important: ______________________________

device before it overheats. When the thermal protector shuts the device off,

## please:

a. Switch the device off.

b. Unplug the device from the electrical outlet.

c. Wait 30 minutes for the motor to cool down before another treatment. Make

sure the air openings are not obstructed.

## Warnings:

. When treating babies, children or anyone who requires assistance, the

following must be observed without fail. Babies, children or anyone who

requires assistance must be supervised constantly by an adult during inhalation

therapy. This is the only way to ensure safe and effective treatment. In addition,

keep the device out of the reach of those that require supervision such as babies

and children.

• This device is designed only for patients who are conscious and can breathe

spontaneously. This is to the only way to ensure effective treatment and avoid a

choking hazard.

For type, dose and regime of medication, follow the instructions of your

doctor or respiratory therapist.

⚫ Clean and disinfect the nebulizer, mouthpiece and mask before using them for

the first time after purchase, if the device has not been used for a long period of

time, or if more than one person uses the device.

• Always rinse the nebulizer, mouthpiece and mask with clean warm tap water

after disinfecting them.

• Never leave the cleaning solution in the nebulizer or mouthpiece..

• Do not store the air tube while there is moisture or medication remaining inside it.

## • Do not use or store the device where it may be exposed to noxious fumes,

flammable gas or volatile substances.

Do not block the air openings. Never place the device where the air openings

may be obstructed during operation. Do not cover the compressor with a blanket

or towel during use.

⚫ Make sure that the nebulizer, mouthpiece and air tube are clean before use.

•Always dispose of any remaining medication after use. Use fresh medication

each time.

Risk of electrical shock

• Never unplug the device with wet hands.

• The compressor and power cord are not waterproof. Do not spill water or

## other liquids, on these parts. If liquid does spill on these parts, immediately

unplug the device and wipe off the liquid.

• Do not use or store the device in humid locations such as a bathroom.

• Do not operate the device with a damaged cord or plug.

• Keep the power cord away from heated surfaces.

## Caution:

## • Make sure that the parts including the filter are attached correctly.

Make sure that the filter is clean. Replace the filter,if the filter has turned

dusty/grey.

Do not tilt the nebulizer kit so the angle of the kit is greater than 45°.

Medication may flow into the mouth. Do not shake the nebulizer kit while using

the device.

• Do not leave the device unattended with babies, children or anyone who

cannot express their consent.

## • Do not subject the device or the parts to any strong shocks such as dropping

the device on the floor.

• Do not disassemble or attempt to repair the compressor or power cord.

## General Safety Precautions:

## Inspect the device and parts before each use for normal operation. Particularly,

## be sure to check the following:

That the nebulizer or air tube is not damaged.

. That the cone of the nebulizer is not blocked.

⚫ That the compressor operates normally. The device generates some noise

and vibration caused by the compressor pump, or some noise caused by the

emission of compressed air from the nebulizer. This is normal and does not

indicate a malfunction.

• Operate the device only as intended. Do not use the device for any other

purpose.

• Do not use the device at temperatures greater than +40°C. Make sure that

the air tube is securely attached to the compressor and nebulizer, and does

not come loose. Twist the air tube plug slightly when inserting it into the

connectors to avoid the tube becoming disconnected during use.

## Disinfecting the nebulizer:

After the last treatment of the day, follow the procedure below to disinfect your

Compressor Nebulizer unless otherwise specified by your doctor

# 1. Gently open the nebulizer.

# 2. Wash the nebulizer, mouthpiece or mask in one part vinegar with three part

## distilled water.Make sure solution is enough to submerge into all part nebulizer

kit. Rinse them thoroughly with solution.

3. Dry them with a clean, soft cloth or let them air dry in a clean environment.

# 4. Reassemble the nebulizer when completely dry. Place these parts in a dry,

sealed container.

## Changing the filter:

1. Do not use cotton or any other material .Do not wash or clean the filter. Use

only filter supplied by distributer. Do not operate without a filter.

# 2. Change the filter when filter turns grey.

3. Changing Procedure: ______________________________

one. Put the filter cap back

## Specifications:

Input voltage

## Current

## Power

## Medication Capacity

## Particle Size

# MMAD

## Sound Level

## Average Nebulizer Rate

## Compressor Pressure Range

## Operating Pressure Range

## Liter Flow Range

## Storage Conditions

## Weight

Shelf-life of Nebulizer kit

## Warranty

AC 220V, 50Hz

# 0.5A

## 45 watt

8ml

0.5 to 10μm

4μm

Less than 60dB at 1 meter

Min. 0.2ml/min.

## 32 Psi to 45 Psi

## 16 to 22 Psi

6-9 L/min

Keep it in a cool and dry place

## 1.4Kg

## 5 years

## 1 Year

# PASSIM LIFESCIENCES LIMITED

## Plot No. 45, Industrial Area Phase 2, Panchkula, Haryana - 134113

Certification of Analysis

Document Title Certification of Analysis for RM/PM/IP/FP

## Revision No.

# R-02

## Effective Date

___/___/_____

## Next Review Date

___/___/_____

## Page No.

Page 1 of 2

Name of the

## Sample:

## Piston Nebulizer

# ONGIO

STS Number: ______________

Date of

___/___/_____

STP Number: ______________

## Sampling:

## Sampling

## Sopes.

## Analytical

## Quantity:

Report (AR)

# OC/FP-043

## No.:

## Material Code:

## Lot No.:

# ONG101

# ONG 23002

Date of

## Analysis:

## Start Date:

## End Date:

___/___/_____ ___/___/_____

## 1.0 Test Report of RM/PM/IP/FP:

## S. No.

## Parameter

## Observed

## Compliance

Done by

## Reading

## Yes

## No

## (Sign &

(✓)

(✓)

## Date)

1.

frequency (Hz)

50

2.

## Voltage (V)

230

# 3. flow rate (4/min)

0.4

# 4. Pressure Range (PST)

45

5.

## Product

## Check Carton

## Justruction

Format: _______________

# MASTER COPY

manual,

## Serial No.

Sign (QA)

## Date:

# ISSUED BY QA

# C-173

___/___/_____

## Biborayo

___/___/_____

## Bikesan

___/___/_____

## Biksany

_____-2

## Bilray

04,02-20

## Bilerany

______

# PASSIM LIFESCIENCES LIMITED

## Plot No. 45, Industrial Area Phase 2, Panchkula, Haryana - 134113

Certification of Analysis

Document Title Certification of Analysis for RM/PM/IP/FP

## Revision No.

# R-02

## Effective Date

___/___/_____

## 2.0 Remark:

## Rass

## Approved by:

## (Sign & Date)

Format: _______________

___/___/_____

## Next Review Date

## Page No.

Page 2 of 2

___/___/_____

## Printing, Packs

ging material ✓

# NA-

# MASTER COPY

## Serial No.

Sign (QA)

## Date:

# ISSUED BY QA

# C-173

___/___/_____

## Bileray

_____-____

## Format No.

## Model

riot No. 45, Industrial Area Phase 2, Panchkula, Haryana - ______

# ONW102

Lot no. _______

## Next Review

___/___/_____

## Date

# F/060/016-00

## Effective Date

___/___/_____

# DESIGN VALIDATION RECORD

Test report for Piston Nebulizer

Lot size

## Date

05

Parameter to be measured

___/___/_____

## Sr. No.

## Product Sr.

Pressure (psi); >30psi

Flow rate(Litre/min;>6)

## Remarks

## Tested By

## No.

(Pass/Fail/NA)

01

_________

## 32 psi

## 8 l/min

## Pass

## Auril

02

## 2206 00002

## 34 Psi

## Slimen

## Pass

## Alenil

03

856

_________

## 36 Psi

81/mon

## Pass

## Alexis

04

_________

## 38 Psi

______

pass

## Alehi

05

_________

## 40 Psi

## Pass

## Are

Prepared by

07

## Caif Of

# MASTER COPY

Page no 1 of 1

## Abhil

Checked by

## Serial No.:

Sign (QA)

## Date:

# ISSUED BY QA

# GA-001

## Looney

___/___/_____

## Format No.

# PASSIM LIFESCIENCES LIMITED

## Plot No. 45, Industrial Area Phase 2, Panchkula, Haryana - 134113

## Model

# ONW102

Lot no. ON W22002

# F/060/016-00

## Effective Date

___/___/_____

# DESIGN VALIDATION RECORD

## Next Review

___/___/_____

## Date

Test report for Piston Nebulizer

Lot size

## Date

05

___/___/_____

Parameter to be measured

## Sr. No.

## Product Sr.

Pressure (psi); >30psi Flow rate(Litre/min;>6)

## Remarks

## Tested By

## No.

(Pass/Fail/NA)

01

_________

34

# PSS

81/min

## Pacs

## Akhil

02

_________

32

## Psi

821min

## Pass

## Akhi

03

_________

38

## Psi

81/min

## Pass

## Akhil

04

_________

36

## Ps:

811mm

## Pass

## Aichil

## 05 220600010

40

## Pass

## Akhil

## Psi

Prepared by

# MASTER COPY

## Dart Of

## Alchil

Checked by

Page no 1 of 1

## Serial No.

Sign (QA)

## Date:

# ISSUED BY QA

# GA-002

___/___/_____

## Format No.

# PASSIM LIFESCIENCES LIMITED

## Plot No. 45, Industrial Area Phase 2, Panchkula, Haryana - 134113

## Model

# ON WI02

Lot no. ONW _____

# F/060/016-00

## Effective Date

___/___/_____

# DESIGN VALIDATION RECORD

Test report for Piston Nebulizer

Lot size

## Date

Parameter to be measured

## Next Review

___/___/_____

## Date

## Sr. No.

## Product Sr.

Pressure (psi); >30psi Flow rate(Litre/min;>6)

## No.

## Remarks

(Pass/Fail/NA)

## Tested By

01

_________

# 36 PS

_____

## Pass

## Akail

02

_________

32

## Psi

81/mm

## Paes

## Achil

03

_________

# 34 PSF

## Pass

## Akhi

04

_________

## 40 Rs

11/mm

pacs

## Alch

## 05 220600015

# 38 PSE

81/mi

## Pass

## Architi

Prepared by

# MASTER COPY

## Cary Of

Page no 1 of 1

## Alchil

Checked by

## Serial No.:

Sign (QA)

## Date:

# ISSUED BY QA

# A003

t

___/___/_____

## Format No.

# PASSIM LIFESCIENCES LIMITED

## Plot No. 45, Industrial Area Phase 2, Panchkula, Haryana-134113

## Model:

## Lot no.:

# 0NW102

# ONW 22001

Parameter to be checked

# F/060/08-00

## Effective Date

___/___/_____

# DESIGN VERIFICATION RECORD

Final Visual Inspection Report for Piston Nebulizer

## Lot size:

05

## Date:

## Next Review

___/___/_____

## Date

___/___/_____

## Product Sr.

## Check

## Check

no.

for

switch

## Scratch/

on/off

## Check

filter

cap

## Check

feet

## Check

all

## Check

## Check

## Check

## Remarks

## Product

medication

user

## Checked

## By

screw

label

## kit all part

manual

dent/

body

fitting

✓

_________

_________

## 2206 00003

_________

_________

Prepared by

broken

✓

✓

✓

Casy ext

_____

# L

☑

Page no 1 of 1

# MASTER COPY

2

✓

# L

# OK

## Akhil

OK Akhil

# OK

# OK

## Alen

## Alen

OK Aren

## Akhil

Checked by

## Akhil

## Serial No.:

Sign (QA)

## Date:

# ISSUED BY QA

# A-001

## Lorey

___/___/_____

## Format No.

# PASSIM LIFESCIENCES LIMITED

## Plot No. 45, Industrial Area Phase 2, Panchkula, Haryana-134113

# F/060/08-00

## Effective Date

___/___/_____

# DESIGN VERIFICATION RECORD

## Model:

## Lot no.:

# ONW102

# ONW 22002

Parameter to be checked

Final Visual Inspection Report for Piston Nebulizer

## Lot size:

05

## Next Review

___/___/_____

## Date

## Date: 28-06-2022

## Check

## Remarks

user

## Checked

## By

manual

## Product Sr.

## Check

## Check

no.

for

switch

## Scratch/

on/off

## Check

filter

cap

## Check

feet

## Check

all

## Check

## Product

screw

label

dent/

## Check

medication

## kit all part

body

broken

fitting

## 22060000 6

_________

✓

✓

✓

✓

_________

_________

£206._____

✓

# D

✓

✓

# L

✓

✓

✓

✓

✓

Prepared by

## Dart Dy

Page no 1 of 1

# MASTER COPY

OK Akhil

# OKA

OK Akhil

OK Aen

OK How

## Arh

Checked by

## Archit

## Serial No.

Sign (QA)

## Date:

# ISSUED BY QA

# 2A-002

___/___/_____

## Format No.

# PASSIM LIFESCIENCES LIMITED

## Plot No. 45, Industrial Area Phase 2, Panchkula, Haryana - 134113

# F/060/08-00

## Effective Date

___/___/_____

## Model:

## Lot no.:

# 0NW102

# ONW 22003

Parameter to be checked

# DESIGN VERIFICATION RECORD

Final Visual Inspection Report for Piston Nebulizer

Lot size: ________

## Date:

## Product Sr.

## Check

## Check

no.

for

switch

## Scratch/

on/off

## Check

filter

cap

## Check

feet

## Check

all

## Check

## Product

screw

label

dent/

broken

body

fitting

_________

_________

_________

_________

_________

Prepared by

Don't Ort

✓

✓

Page no 1 of 1

# MASTER COPY

## Next Review

___/___/_____

## Date

___/___/_____

## Check

## Remarks Checked

## By

## Check

medication

## kit all part

user

manual

# SS SSS

# OK

## Alent

# OK

## Alehil

ok

## Ales

# OK

## Allis

OK Arch's

Checked by

## Alzhis.

20.06.____

## Serial No.

## Sign (QA):

# ISSUED BY QA

# QA-003

## Son

## Batch Manufacturing Record- BMR

Format No. PLL/BMR/01-03

## Effective Date

## 13-06-2024 Next Review Date 12-06-2029

## BMR Document No.:

## MFR No.:

## Product Generic Name:

## Model No.:

## Lot/Batch No.:

## Lot Size:

# PLYMFRI 006

## Piston Nebulizer

# ONS 105

# ONS 105 24009

## 2070 pcs.

## Product Serial No.:

## Mfg. Date:

## Storage condition:

## Date of Batch Initiation:

## Date of Batch Completion:

## Shelf life:

___________

___________

## To

## July-2024

## Keep

it in a cool and dry placs.

___/___/_____

___/___/_____

-

# NA

-

Page 1 of 7

# MASTER COPY

## Serial No.

# ISSUED BY QA

# PR-012

Sign (QA)

## Date:

13-____

## Batch Manufacturing Record- BMR

Format No. PLL/BMR/01-03

## Effective Date 13-06-2024

## Next Review Date 12-06-2029

## 1.0 Bill Of Raw Material:

Material Requisition Slip (F/043/03-current revision) is attached.

## Material Code

## Description

## Standard Quantity Total Quantity

per product

as per Lot

## Rm595

## Power Cord

2070x1

## Rm 182

## Rm 183

# RM 597

# RM185

## Back Stopper Rubber

Air filter round shape

Luutch on/off

4

______

1

# 2070 XI

1

# 2070 XI

## Silicon Tube

## 400 mts

## 400 mtr

# RM18

# RM 187

Cable tie

## Ss. Spring

2

2070x2

3

# RM 188

Sonewsph Type sk

6x13

4

# 2070 X3

2070x4

# RM 189

Lubber pad

4

2070x4

# RM 190

Medication kit

1

# 2070X1

# RM 191

## Compressor Motor

2070x1

# RM312

## Top Cabinet

2070x1

# RM603

## Bottom Cabinet

1

# 2070 XI

# RM606

Air Filter Cap blue

# 2070XI

## Rm 195

Rubber band

·500kg

0.500kg

# RM 594

Glue stick

## 6 plat

## 6Pkt

# RM 370

## Insulation Tape

# 12 PC

12pc

# NA

Page 2 of 7

# MASTER COPY

## Serial No.

# ISSUED BY QA

# PR-012

Sign (QA)

## Date:

___/___/_____

## Batch Manufacturing Record- BMR

Format No. PLL/BMR/01-03

## Effective Date 13-06-2024 Next Review Date

___/___/_____

## 2.0 Bill Of Packing Material:

Material Requisition Slip (F/043/03-current revision) is attached.

## Standard Quantity

## Material Code

## Total Quantity

## Description

per product

# PM 496

## Instruction Manual

1

## Pm124

## Pm125

## Pm 335

pm224

pm 233

# PM247

Polybag 5x7

Non wouch bag

## Carton

Comugated box

## Surface Cream

Air Filter round shape

## (Accessories)

## 6 kg

as per Lot

# 2070X1

61g

# 2070 XI

# 2070X1

20___/___/_____

## 173 boxes

10

## Pcs

1

# 2070 XI

## 3.0 Assembly Instructions:

# NA

The quality control person shall provide the line clearance by checking the below parameter

## to initiate the batch manufacturing:

3.1) Verify the cleaning status of area and ensure all the material (Raw, in-process) and

packed finished product have been removed from the line/area.

3.2) Verify & ensure that assembly line/ production area are properly cleaned.

Page 3 of 7

# MASTER COPY

# L

## Serial No.

Sign (QA)

## Date:

# ISSUED BY QA

# PR-012

___/___/_____

## Batch Manufacturing Record- BMR

Format No. PL/BMR/01 0. Effective Date

## 2024 Next Review Date

12.06 ____

## Line Clearance For Manufacturing Process

Manufacturing/Packing area cleaned and free from materials, containers of previous product

## (Yes/No):

## Previous Product

# NA

Area cleaned as per SOP No. PLL/SOP/041

## Checked By / Date:

## (Production)

____

## Batch/Lot No

## Done By /Date

# NA

# THE

## Certified By /Date:

## Bitter

# (QC)

03-______

## 4.0 Assembly:

## Assembly Start

## S. No.

## Date

## Time

## Date

## Assembly End

## Time

No. of Product

## Done Checked

## Assembled

by

by

1

03-08-2024 09: ________

03-08-2024 18: ________

## 517 PCs.

## Indeep Achi

## 2 05-08-2024 09:30

3

05-08-2024 18: ________

06-08-2024 09: _______________

## 517 pus

smdlip Akki

## 518 p

4

07-18-2024 09: _______________

# 518 ICS

## Sondery Desi

## Sundeep Aktu

5

6

7

8

9

10

11

12

13

14

15

# VA

## 5.0 Packing Instruction:

The quality control person shall provide the line clearance by checking the below parameter

to initiate the packing of the assembled product: _______________

# MASTER COPY

Page 4 of 7

## Serial No.

Sigu (QA)

## Date:

# ISSUED BY QA

____

___/___/_____

## Batch Manufacturing Record- BMR

Format No. PLL/BMR/01-03

## Effective Date 13-06-2024

## Next Review Date 12 06-2029

5.1) Verify the cleaning status of area and ensure all sort of unwanted packing material have

been removed from the line/area.

5.2) Verify & ensure that packing line is properly cleaned and dust free.

## Line Clearance For Packing Process

Manufacturing/Packing area cleaned and free from materials, containers of previous product

## (Yes/No):

## Previous Product:

# NA

Area cleaned as per SOP No. PLL/SOP/041

## Checked By / Date: ....

## (Production)

___/___/_____....

## 6.0 Packing:

## Batch/ Lot No:

# NA

## Done By /Date:

## Hidrur

Certified By /Date: ___/___/_____

# (QC)

Packing start

Packing end

No. of product

## Done Checked

## S. No.

packed

by

by

## Date

## Time

## Date

## Time

1

03-08-2024 09: _______________

18: ________

## 517 Pcs fondeep Aktuel

2

___/___/_____ _____

05-08.

8.____

18: ________

## 517 pcs

## Sanders Actu

3

06-08-2024 09: _______________

18: ________

## 518 pcs

4

___/___/_____

09: _______________

18: ________

## Sundeep Akku

## 518 pes' Sendup Actul

5

6

78

9

10

## 7.0 QC Sample Quantity:

# 21PC3-

(CoA copy is attached with the BMR)

8.0 Transfer Finished Goods to Finished Goods Store

## QC release Status:

foleased...

(Transfer Ticket - F/044/01-current revision is attached with the BMR)

# MASTER COPY

Page 5 of 7

## Serial No.

# ISSUED BY QA

# PR-012

Sign (QA)

## Date:

___/___/_____

## Batch Manufacturing Record- BMR

Format No. PLL/BMR/01-03

## Effective Date 13-06-2024 Next Review Date

## 9.0 Reconcilation Of Material:

## S. No. Material

## Code

1

2

3

4

5

6

7

8

9

___/___/_____

## Description

## Remaining

## Quantity

## Action Taken

## Reference

## Record

# MA/RN

## Attached

10

Note: ______________________________

any).

100%

## 10.0 Theoretical Yield:

100%

## 11.0 Actual Yield:

# NIL

## 12.0 Rejection:

## 13.0 Control Sample:

## 02 pcs.

## 14.0 Date of Completion:

___/___/_____

15.0 AR No. /Final test report number and Date: ___/___/_____

Page 6 of 7

# MASTER COPY

## Serial No.

Sign (QA)

## Date:

# ISSUED BY QA

# PR012

___/___/_____

## Batch Manufacturing Record- BMR

Format No. PLL/BMR/01-03

## Effective Date 13-06-2024 Next Review Date

___/___/_____

## 16.0 Attachment of Labels:

(A4 sheet to be attached in case of shortage of space for labels)

## Kushal

## Kushal

## Prepared By

## (Production Officer)

# MASTER COPY

## Dr. Odin

## Piston Nebulizer

# ONS105

## Voltage: 220V

# LOT ONS10524009

# SN 24070900001

07-____

# MDA

## Current: 0.5A

# ISO

____

## Passim Lifesciences Ltd.

## Plot No. 48, Ind. Area, Phase-11,

Panchkula-______ (HR)-INDIA

## Mfg. Lic. No.

# MPO/MD/2012/000633

Email ID: ________@_______.___

Website: _______________

## Actui

## Verified By

## (Production Incharge)

Page 7 of 7

kshut

## Verified By

## (Quality Incharge)

## Serial No.

Sign (QA)

## Date:

# ISSUED BY QA

# PR-012

___/___/_____

## Batch Manufacturing Record- BMR

Format No. PLL/BMR/01-03

## Effective Date 13-06-2024 Next Review Date 12-06-2029

## BMR Document No.:

## MFR No.:

## Product Generic Name:

## Model No.:

## Lot/Batch No.:

## Lot Size:

## Product Serial No.:

## Mfg. Date:

## Storage condition:

## Date of Batch Initiation:

## Date of Batch Completion:

## Shelf life:

# PLL BMR ONS 24-075

# PLL MFR 006

## Piston Nebulizer

# ONS 105

# 6N510524010

## 2020 pcs.

## 24071000001 To 24071002020

## July-2024

Keep it in a cool and dry place

___/___/_____

___/___/_____

# NA-

Page 1 of 7

# MASTER COPY

# ISSUED BY QA

Serial No. PR-028

Sign (QA)

## Date: 01-09-2024

## Batch Manufacturing Record- BMR

Format No. PLL/BMR/01-03 Effective Date ___/___/_____ Next Review Date ___/___/_____

## 1.0 Bill Of Raw Material:

Material Requisition Slip (F/043/03-current revision) is attached.

## Material Code

## Description

## Standard Quantity Total Quantity

per product

as per Lot

# RM595

## Power Cord

1

# RM182

## Back Stopper Rubber

4

# RM183

# RM 597

Air filter round shope

# 2020 XI

2020x4

# 2020 XI

witch on/off

1

# 2020 XI

# RM 185

## Silicon Tube

## 400 mtr

## 400 mtr

# RM186

# RM 187

S.s.

Cable fring

2

# 2020 X2

3

# 2020X3

# RM 188

Screw Ph Jupe 6x13

4

# 2020X4

# RM 189

Rubber fad

4

20.____

# RM 190

Medication kit

# 2020X1

## Rm 191

# RM 312

## Top Cabinet

# AM 603

## Compressor Motor

## Bottom Cabinet

# 2020X1

# 2020X1

# 2020X1

# RM606

## Air Filter

## Cop

## Blue

# 2020X1

# RM 195

Rubber band

0'

• 500kg

0.500kg

# RM594

## Glue Stick

## 6 pkt

## 6 pret

## Rm 370

## Insulation Tape

10

# 10PG

Page 2 of 7

# MASTER COPY

# ISSUED BY QA

## Serial No.

# PR-028

Sign (QA)

## Date: 01-09-2024.

## Batch Manufacturing Record- BMR

## Format No.

# PLL/BMR/01-03

## 2.0 Bill Of Packing Material:

## Effective Date 13-06-2024 Next Review Date 12-06-2029

Material Requisition Slip (F/043/03 current revision) is attached.

## Material Code

# PM 496

# PM124

# PM 125

pm 335

pm224

pm233

## Pm247

## Description

## Instruction Manual

## Polybag 5X7

Non woven bag

## Carton

Comugated box

## Surface Cream

Air filter round shape

## Standard Quantity Total Quantity

per product

1

1

1

kg

____+12

# 10 PC

as per Lot

# 2020X1

6kg

# 2020X1

# 2020X1

## 169 pes

Lo pe

# 2020 X1

## Accessories

# NA

## 3.0 Assembly Instructions:

The quality control person shall provide the line clearance by checking the below parameter

## to initiate the batch manufacturing:

3.1) Verify the cleaning status of area and ensure all the material (Raw, in-process) and

packed finished product have been removed from the line/area.

3.2) Verify & ensure that assembly line/ production area are properly cleaned.

Page 3 of 7

# MASTER COPY

# ISSUED BY QA

## Serial No.

# PB-028

## Sign (QA):

## Date: 01-09-2024

## Batch Manufacturing Record- BMR

Format No. PLL/BMR/01-03 Effective Date ___/___/_____ Next Review Date

___/___/_____

## Line Clearance For Manufacturing Process

Manufacturing/Packing area cleaned and free from materials, containers of previous product

## (Yes/No):

## Previous Product:

# NA

## Checked By Date:

Area cleaned as per SOP No. PLL/SOP/041

## (Production)

## Nieve

___/___/_____

## Batch/ Lot No:

# NA

## Done By /Date:

## Certified By /Date:(

# (QC)

________..

## Susat

## 4.0 Assembly:

## Assembly Start

## Assembly End

No. of Product

## Done

## S. No.

## Assembled

by

## Checked

by

## Date

## Time

## Date

## Time

1

25-09-2014 09: _______________

## 674 pcs

2

3

26-09-2024 09: _______________

27-09-2014 09: _______________

## 674 Pcs

18: _______________

## Vikas Aktu

## Vikas Akhil

## Vikas Aptu

4

10

5

6

# IN

7

8

9

____

11

# NA

15

14

15

## 5.0 Packing Instruction:

The quality control person shall provide the line clearance by checking the below parameter

to initiate the packing of the assembled product: _______________

# MASTER COPY

Page 4 of 7

# ISSUED BY QA

## Serial No. R-028

Sign (QA)

___/___/_____

## Batch Manufacturing Record- BMR

Format No. PLL/BMR/01-03 Effective Date ___/___/_____ Next Review Date

___/___/_____

5.1) Verify the cleaning status of area and ensure all sort of unwanted packing material have

been removed from the line/area.

5.2) Verify & ensure that packing line is properly cleaned and dust free.

## Line Clearance For Packing Process

Manufacturing/Packing area cleaned and free from materials, containers of previous product

## (Yes/No):

## Previous Product:

# NA

## Checked By Date: ......

Area cleaned as per SOP No. PLL/SOP/041

## (Production)

## Nevio

________

## Batch/ Lot No:

# NA

## Done By /Date:

## Certified By /Date:

# (QC)

tcy

## Date: Suray

________

## 6.0 Packing:

Packing start

Packing end

## S. No.

No. of product

packed

## Done Checked

by

by

## Date

## Time

## Date

## Time

1

25.09.2024 09: _______________

_____

2

26-09-2024 09: ________

___/___/_____

18: ________

3

27-09-2024 09: _______________

18: ________

## 674 pcs Vikas Adic

## 674 Pcs Vikas Aktu

## 672 pcs. Vikal. Alchi

4.

5

67

00

8

9

10

10

# ·NA

7.0 QC Sample Quantity: ______

(CoA copy is attached with the BMR)

8.0 Transfer Finished Goods to Finished Goods Store

QC release Status: ________

## Released

(Transfer Ticket - F/044/01-current revision is attached with the BMR)

# MASTER COPY

Page 5 of 7

# ISSUED BY QA

Serial No. PR-028

Sign (QA)

## Photo 01-09-2024

## Batch Manufacturing Record- BMR

Format No. PLL/BMR/01-03

## 9.0 Reconcilation Of Material:

## S. No.

## Material

## Code

1

23

## Effective Date

## 13-06-2024 Next Review Date

___/___/_____

## Description

## Remaining

## Action Taken

## Quantity

## Reference

## Record

4.

567

8

9

10

# MRN

## Attached

Note: ______________________________

any).

## 10.0 Theoretical Yield:

30.

## 11.0 Actual Yield: .

100%

## 12.0 Rejection:

# NIL

## 13.0 Control Sample:

02

## Pes

___/___/_____

## 14.0 Date of Completion:

15.0 AR No. /Final test report number and Date: _______________

Page 6 of 7

# MASTER COPY

# 24/PP/24-0106

# ISSUED BY QA

Serial No. PR-028

Sign (QA)

## Date: 01-09-2024

## Batch Manufacturing Record- BMR

## Effective Date 13-06-2024

## Next Review Date

___/___/_____

## Format No.

# PLL/BMR/01-03

## 16.0 Attachment of Labels:

(A4 sheet to be attached in case of shortage of space for labels)

## Kishal

## Dr. Odin

## Piston Nebulizer

# ONS105

## Voltage: 220V

# LOT ONS10524010

# SN 24071000001

07-____

# MDA

## Current: 0.5A

45

## Passim Lifesciences Ltd.

## Plot No. 45, Ind. Area, Phase-11,

Panchkula-______ (10)- INDIA

## My Lie. Na

# MFG/MD/2022/000433

Email ID: ________@_______.___

Website: _______________

## Nevroy

## Prepared By

## (Production Officer)

## Verified By

## (Production Incharge)

# MASTER COPY

Page 7 of 7

## Verified By

## (Quality Incharge)

# ISSUED BY QA

Serial No.: ________

## Sign (04)

## Date: 01-09-2024

## Batch Manufacturing Record- BMR

Format No. PLL/BMR/01-03

## Effective Date

## 13-06-2024 Next Review Date

12: ________

## BMR Document No.:

## MFR No.:

## Product Generic Name:

## Model No.:

## Lot/Batch No.:

## Lot Size:

## Product Serial No.:

## Mfg. Date:

## Storage condition:

## Date of Batch Initiation:

## Date of Batch Completion:

## Shelf life:

# PLL MFR 006

## Piston Nebulizer.

# ONS105

# ONS 10524011

## 2030 pcs.

___________

___________

## To

## July-2024

keep it in a cool and dry place.

___/___/_____

___/___/_____

# NA-

Page 1 of 7

# MASTER COPY

# ISSUED BY QA

## Bedal No

Sign (QA)

## Date:

# PR-023

___/___/_____

## Batch Manufacturing Record- BMR

Format No. PLL/BMR/01-03

## 1.0 Bill Of Raw Material:

## Effective Date 13-06-2024 Next Review Date

___/___/_____

Material Requisition Slip (F/043/03-current revision) is attached.

## Standard Quantity Total Quantity

## Material Code

## Description

# RM 595

## Power Cord

# RM 182

## Bm 183

# RM 591

# RM 185

# RM 186

## Back Stopper Rubber.

An Filter round shape

on/off switch

## Silicon Tube

## S.S. Specing

per product

1

4

as per Lot

# 2030X1

# 2030 X4

# 2030 X1

1

# 2030X1

## 400 Mtr

## 400 mt

2

# 2030X2

## Rm187

## Cable Te

3

# 2030X3

# RM188

Screw PhType 6x13

4

# 2030 X4-

# RM 189

## Rubber Pad

4

______

# RM 190

Medication kit

2030x1

## Rm191

## Compressor Motor

2030x1

# RM312

## Top Cabinet

2030x1

## Rm602

## Bottom Cabomit

1

# RM605

## Amr Filter Cap

2030x1

# 2030X1

## Rm 195

Rubber band

## Rm594

Glue stick

0.500 kg

# 6PKJ

0.500kg

## 6 Pkt

# RM370

## Insulation Tape

## 10 Pe

## 10 pc

# NA

Page 2 of 7

# MASTER COPY

## Serial Net

Sign (QA)

## Date:

# 19SUED BY QA

# PR-023

___/___/_____

## Batch Manufacturing Record- BMR

Format No. PLL/BMR/01-03

## Effective Date 13-06-2024 Next Review Date 12-06-2029

## 2.0 Bill Of Packing Material:

Material Requisition Slip (F/043/03-current revision) is attached.

## Material Code

## Description

## Standard Quantity

per product

## Total Quantity

as per Lot

## Pm498

## Instruction Manual

# 2030X1

pm124

poly bag 5x7

## 6 kg

6kg

# M125

No won

## Pm414

pm224

# PM 223

# PM247

neoven bag

## Carton

Comugated box

## Surface Cream

Air filter round shape.

## Accessory.

______

# 10 PC

# 2830X1

# 2030X1

## 170 boxes

10pc

# 2030X)

# NA

## 3.0 Assembly Instructions:

The quality control person shall provide the line clearance by checking the below parameter

## to initiate the batch manufacturing:

3.1) Verify the cleaning status of area and ensure all the material (Raw, in-process) and

packed finished product have been removed from the line/area.

3.2) Verify & ensure that assembly line/ production area are properly cleaned.

Page 3 of 7

# MASTER COPY

## Serial Not

## Sign (0)

## Date:

# ISSE FILIY QA

# PR-023

# F

___/___/_____

## Batch Manufacturing Record- BMR

Format No. PLL/BMR/01-03

## Effective Date 13-06-2024 Next Review Date

## Line Clearance For Manufacturing Process

___/___/_____

Manufacturing/Packing area cleaned and free from materials, containers of previous product

## (Yes/No):

## Previous Product:

# NA

Area cleaned as per SOP No. PLL/SOP/041

## Batch/ Lot No: ....

## Done By /Date:

## Checked By / Date: ......

.....

## (Production)

07-09 ____.....

# NA

## Certified By /Date: May.

# (QC)

___/___/_____".

## 4.0 Assembly:

## Assembly Start

## Assembly End

No. of Product

## Done

## Checked

## S. No.

## Assembled

by

by

## Date

## Time

## Date

## Time

1

07-09-2024 09: ________

2

09.09.____

09: ________

09-09-2024 18: ________

## 1015 pcs. Curpreet Aut

## 1015 pcs

## Current Delle

3

4

5

67

## 8 9

10

11

12

13

14

15

## 5.0 Packing Instruction:

The quality control person shall provide the line clearance by checking the below parameter

to initiate the packing of the assembled product: _______________

# MASTER COPY

Page 4 of 7

## Serial No

Sign (QA

## Date:

# ISSUED BY QA

023

f

___/___/_____

## Batch Manufacturing Record- BMR

Format No. PLL/BMR/01-03

## Effective Date 13-06-2024 Next Review Date 12-06-2029

5.1) Verify the cleaning status of area and ensure all sort of unwanted packing material have

been removed from the line/area.

5.2) Verify & ensure that packing line is properly cleaned and dust free.

## Line Clearance For Packing Process

Manufacturing/Packing area cleaned and free from materials, containers of previous product

## (Yes/No):

## Previous Product:

# NA

Area cleaned as per SOP No. PLL/SOP/041

## Checked By / Date:......

## (Production)

## .. Kusshay

17.09.____.

# NA

## Batch/Lot No:

## Done By /Date:

Certified By /Date: ___/___/_____

# (QC)

___/___/_____

## 6.0 Packing:

Packing start

Packing end

## S. No.

No. of product

packed

## Done Checked

by

by

## Date

## Time

1

07-______ _____

## Date

___/___/_____

## Time

18: ________

2

09-09-2024 09: _______________

## 1015 Pcs.

## Current Achi

## 1015 pcs. Current Acrui

3

4

5

6

7

8

9

10

7.0 QC Sample Quantity: __________________

(CoA copy is attached with the BMR)

8.0 Transfer Finished Goods to Finished Goods Store

## ... Released....

QC release Status: ________

(Transfer Ticket - F/044/01-current revision is attached with the BMR)

# MASTER COPY

Page 5 of 7

## Serial Nost

## Sign

## Date:

# ESSED BY O

# PR-023

___/___/_____

## Batch Manufacturing Record- BMR

Format No. PLL/BMR/01-03 Effective Date

___/___/_____

## 9.0 Reconcilation Of Material:

## S. No.

## Material

## Code

1

2

3

4

5

6

7

8

9

## Next Review Date 12-06-2029

## Description

## Remaining

## Action Taken

## Quantity

## Reference

## Record

# NO

10

Note: ______________________________

any).

100%

## 10.0 Theoretical Yield:

## 11.0 Actual Yield:

100%

# NIL

## 12.0 Rejection:

## 13.0 Control Sample:

-___/___/_____ 02 pcs.

## 14.0 Date of Completion:

___/___/_____

15.0 AR No. /Final test report number and Date: _______________

Page 6 of 7

# MASTER COPY

# QC/FP/24-0101

## Serial

## Sign (0)

## Dat

# PR-023

___/___/_____

## Batch Manufacturing Record- BMR

Format No. PLL/BMR/01-03 Effective Date

## 13-06-2024 Next Review Date 12-06-2029

## 16.0 Attachment of Labels:

(A4 sheet to be attached in case of shortage of space for labels)

## Kesha

## Dr. Odin

## Piston Nebulizer

# ONS105

## Voltage:220V Current: 0.5A 145

# LOT ONS10524011

# SN 24071100001

07-____

# MDA

## Passim Lifesciences Ltd.

## Plot No. 45, Ind. Area, Phase-It,

Panchkula-______ (HR)-INDIA

## Mly. Lic. No.:

# MFG/MD/2011/000633

Email ID: ________@_______.___

Website: _______________

## Kehal

## Prepared By

## (Production Officer)

# MASTER COPY

## Aemi

## Verified By

## (Production Incharge)

Page 7 of 7

afshet

## Verified By

## (Quality Incharge)

## Serial

## Sign (0

## Date:

# LMED BY OA

## Ph023

f

___/___/_____

================================================================================

## 📋 INTELLIGENT TEMPLATE INSTRUCTIONS

🧠 This template was created using intelligent content analysis

📝 All fields were automatically identified and converted to appropriate blanks

📊 Original table structure and formatting has been preserved

🎯 Fill in all underlined blank fields with appropriate information

✅ No manual pattern matching was used - fully dynamic analysis

💡 Field Types Detected:

• ___/___/_____ = Date fields

• ___@___.__ = Email addresses

• ___-___-____ = Phone numbers

• _______ = General text fields

🔧 If any field appears incorrect, please adjust as needed.